NP Photonics, Inc.

U of A Science and Technology Park

9030 S. Rita Road, Suite 120

Tucson, AZ 85747

Telephone: 520-799-7400

Fax: 520-799-7403

www.NPPhotonics.com

Fiber Laser Controller

Command Language Specification

| Document Number | 690001 (FLM-005) | Document Classification | INTERNAL & CUSTOMER USE |
| --- | --- | --- | --- |

| Issue | Date | ECN | Written By | Authorised By |
| --- | --- | --- | --- | --- |
| A | 07-Jan-04 | - | CMC |  |
| B | 21-Jan-04 | - | CMC |  |
| C | 11-Feb-04 | - | CMC |  |
| D | 26-Feb-04 | - | CMC |  |
| E | 05-Mar-04 | - | CMC |  |
| F | 22-Mar-04 | - | CMC |  |
| G | 14-Apr-04 | - | CMC |  |
| H | 02-Jun-04 | - | CMC |  |
| I | 17-Aug-04 | - | CMC |  |
| J | 25-Aug-04 | - | CMC |  |
| K | 10-Sep-04 | - | CMC |  |
| L | 11-Oct-04 | - | CMC |  |
| M | 22-Feb-05 | - | CMC |  |
| N | 27-Sep-05 | - | CMC |  |
| O | 07-Oct-05 | - | CMC |  |
| P | 02-Nov-05 | - | CMC |  |
| Q | 27-Jan-06 | - | CMC |  |
| R | 02-Aug-06 | - | CMC |  |
| S | 05-Dec-07 | - | CMC |  |
| T | 06-Feb-08 | - | CMC |  |
| U | 04-Aug-08 | - | CMC |  |
| V | 06-Oct-08 | - | CMC |  |
| W | 16-Mar-09 | - | CMC |  |
| X | 29-Jun-09 | - | CMC |  |
| Y | 16-Mar-10 | - | CMC |  |
| Z | 19-May-10 | - | CMC |  |
| AA | 22-Nov-11 | N00225 | CMC |  |
| AB | 20-Aug-15 | N00284 | CMC |  |
| AC | 05-Jan-17 | N00285 | CMC |  |
| AD | 05-Dec-18 | N00290 | CMC |  |
| AE | 26-Mar-19 | N00292 | CMC |  |
| AF | 02-Jul-19 | N00305 | CMC |  |
| AG | 25-Oct-22 | N00313 | CMC |  |
| AH | 24-Oct-22 |  | JHJ |  |

Table of Contents

Revision History	6

1	Introduction	12

1.1	Related Documents	12

1.2	References	12

2	Hardware Interface	13

2.1	RS-232 Connections – via Interface Module	13

2.2	RS-232 Connections – direct to Laser Module	13

3	Command Language Overview	14

3.1	General	14

3.2	Command Syntax	14

3.2.1	Commands and Command Parameters	14

3.2.2	Query Commands	14

3.2.3	Case Sensitivity	14

3.2.4	Long-Form and Short-Form	15

3.2.5	Program Messages	15

3.2.6	Program Message Terminator	15

3.2.7	Response Messages	15

3.2.8	Response Message Terminator	15

3.2.9	Multi-Channel Addressing	16

3.2.9.1	Physical Addressing	16

3.2.9.2	Aliased Addressing	16

4	Common Commands	17

4.1	*CLS – Clear Status	17

4.2	*IDN? – Identification Query	17

4.3	*OPT? – Option Identification Query	17

4.4	*RST – Reset	17

4.5	*SAV – Save	17

5	SCPI Command Reference	18

5.1	:SYSTem Commands	18

5.1.1	:AMPLifier Commands	19

5.1.1.1	:CONTrol Commands	26

5.1.2	:BFACet[0|1] Commands	27

5.1.3	:COMMunicate Commands	28

5.1.3.1	:ADDRess Commands	28

5.1.3.2	:FORWard Commands	29

5.1.3.3	:SERial<0> Commands	30

5.1.3.4	:SERial<1> Commands	30

5.1.4	:ERRor Commands	32

5.1.5	:FIRMware Commands	32

5.1.6	:HARDware Commands	33

5.1.7	:MHFTuning Commands	36

5.1.7.1	:BAND<a|i> Commands	37

5.1.8	:OPTimization Commands	41

5.1.8.1	:DIP Commands	42

5.1.8.2	:PEAK Commands	43

5.1.8.3	:QUADratic Commands	45

5.1.8.4	:RAMP Commands	46

5.1.9	:PADJust Commands	48

5.1.10	:PASSword Commands	50

5.1.11	:PCONtrol Commands	51

5.1.11.1	:MODE<0|1> Commands	55

5.1.11.2	:RANGe<0[|1|2|3]> Commands	57

5.1.12	:PDRive Commands	57

5.1.12.1	:MODulation Commands	58

5.1.13	:PMONitor Commands	60

5.1.14	:RIN Commands	62

5.1.14.1	:ADJust Commands	64

5.1.15	:RTIMe<0|1> Commands	65

5.1.16	:SMODe Commands	65

5.1.16.1	:CORRection Commands	66

5.1.16.1.1	:RAMP Commands	68

5.1.17	:TCONtrol<0[|1|2|3|4]> Commands	69

5.1.18	:TMONitor<0[|1|2|3|4]> Commands	75

5.1.19	:VOA Commands	76

5.1.19.1	:DATa<i> Commands	77

5.1.19.2	:MODulation Commands	78

5.1.20	:WADJust Commands	79

5.1.21	:Shutdown Commands	82

5.2	:STATus[0|1|2|3] Commands	83

5.3	:MESSage Commands	85

5.4	:DIAGnostic Commands	86

5.4.1	:AMPLifier Commands	86

5.4.2	:TEST Commands	86

5.4.3	:WDOG Commands	87

5.4.3.1	Watchdog Status Report	89

5.5	:OPTimization Commands	89

5.5.1	:LOG Commands	93

5.5.1.1	:EVENt<i> Commands	93

5.5.2	:RAMP<0|1> Commands	94

List of Tables

Table 4.1 – Common Command Summary	17

Table 5.1 – Amplifier Status Flags	21

Table 5.2 – Amplifier Shutdown Codes	21

Table 5.3 – Amplifier Operating Modes	21

Table 5.4 – Amplifier Options	22

Table 5.5 – Amplifier Power Monitors	23

Table 5.6 – Amplifier Tap Factors	25

Table 5.7 – Amplifier Power Thresholds	25

Table 5.8 – Amplifier Temperature Monitors	26

Table 5.9 – Back Facet Flags	27

Table 5.10 – PCB Options	35

Table 5.11 – PCB Types	36

Table 5.12 – Pump Control Flags	52

Table 5.13 – Pump Operating Modes	52

Table 5.14 – RIN Suppression Flags	63

Table 5.15 – RIN Suppression Modes	63

Table 5.16 – RIN Auto-Adjust Flags	64

Table 5.17 – Second Mode Flags	65

Table 5.18 – Second Mode Correction Flags	67

Table 5.19 – Second Mode Correction States	68

Table 5.20 – Temperature Control Channels	69

Table 5.21 – Heater Element Configuration	70

Table 5.22 – Temperature Control Flags	71

Table 5.23 – Temperature Monitor Channels	75

Table 5.24 – Temperature Monitor Flags	75

Table 5.25 – VOA Modulation States	79

Table 5.26 – VOA Modulation Waveforms	79

Table 5.27 – Wavelength Adjust Flags	80

Table 5.28 – Wavelength Adjustment States	81

Table 5.29 – Parameter List returned by :STATus[0|1|2|3]:ALL?	83

Table 5.30 – Parameter List returned by :STATus[0|1|2|3]:FLAGs?	84

Table 5.31 – Optimization Flag Values	89

Table 5.32 – SYSTem Command Summary	95

Table 5.33 – STATus Command Summary	100

Table 5.34 – MESSage Command Summary	100

Table 5.35 – DIAGnostic Command Summary	100

Table 5.36 – OPTimization Command Summary	101

Table 5.37 – Error Messages	101

# Revision History

| Issue | Date | Section | Details |
| --- | --- | --- | --- |
| AH | 24-Oct-22 | 1.1 5.1.21 Table 5.32 | Related documents updated. :System:Shutdown commands added Table updated |
| AG | 25-Oct-22 | 1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.11 5.1.17 5.1.17 Table 5.33 | Related documents updated. Introduction to :SYSTem:AMPLifier commands revised. :SYSTem:AMPLifier:C[0¦1¦2]0 commands revised. :SYSTem:AMPLifier:C[0¦1¦2]1 commands revised. :SYSTem:AMPLifier:CMAXimum<[0]¦1> commands revised. :SYSTem:AMPLifier:CMINimum<[0]¦1> commands revised. :SYSTem:AMPLifier:CSETpoint<[0]¦1> commands revised. :SYSTem:AMPLifier:P[0¦1]0 commands revised. :SYSTem:AMPLifier:P[0¦1]1 commands revised. :SYSTem:AMPLifier:RRATe<[0]¦1> commands revised. :SYSTem:PCONtrol<0¦1>:RRATe commands added. :SYSTem:TCONtrol<1¦2>:OMAX commands revised. :SYSTem:TCONtrol<1¦2>:OMIN commands revised. Table updated. |
| AF | 02-Jul-19 | 1.1 Table 5.6 | Related documents updated. Table updated. |
| AE | 26-Mar-19 | 1.1 5.1.1 5.1.1 Table 5.5 Table 5.33 | Related documents updated. :SYSTem:AMPLifier:OPTions commands added. :SYSTem:AMPLifier:RRATe<[0]¦1> commands revised. Table added. Table updated. |
| AD | 05-Dec-18 | 1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 Table 5.3 Table 5.6 Table 5.8 Table 5.33 | Related documents updated. :SYSTem:AMPLifier:C[0¦1¦2]0 commands revised. :SYSTem:AMPLifier:C[0¦1¦2]1 commands revised. :SYSTem:AMPLifier:CMAXimum<[0]¦1> commands revised. :SYSTem:AMPLifier:CMINimum<[0]¦1> commands revised. :SYSTem:AMPLifier:CMONitor? command revised. :SYSTem:AMPLifier:CSETpoint<[0]¦1> commands revised. :SYSTem:AMPLifier:P[0¦1]0 commands revised. :SYSTem:AMPLifier:P[0¦1]1 commands revised. :SYSTem:AMPLifier:PMAXimum commands revised. :SYSTem:AMPLifier:PMINimum commands revised. :SYSTem:AMPLifier:PMONitor<[0]¦1>? command revised. :SYSTem:AMPLifier:PSETpoint commands revised. Note added for 2-stage Power Rock. Notes added for 2-stage Power Rock. Note added for 2-stage Power Rock. Table updated. |
| AC | 05-Jan-17 | 1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.1 5.1.2 5.1.3.2 5.1.3.4 5.1.6 5.1.7.1 5.1.7.1 5.1.8 5.1.11 5.1.9 5.1.10 5.1.11.1 | Related documents updated. Introduction to :SYSTem:AMPLifier commands revised. :SYSTem:AMPLifier:C[0¦1]0 commands revised. :SYSTem:AMPLifier:C[0¦1]1 commands revised. :SYSTem:AMPLifier:CMONitor? command added. :SYSTem:AMPLifier:FLAGs? command revised. :SYSTem:AMPLifier:P[0¦1]0 commands revised. :SYSTem:AMPLifier:P[0¦1]1 commands revised. :SYSTem:AMPLifier:PMONitor<[0]¦1>? command revised. :SYSTem:AMPLifier:RRATe commands added. :SYSTem:AMPLifier:TFACtor[0¦1] commands revised. :SYSTem:AMPLifier:TMONitor<0¦1>? command added. References to FLM-0177 board added. Reference to FLM-0177 board added. Reference to FLM-0177 board added. References to FLM-0177 board added. Introduction to :SYSTem:MHFTuning:BAND<a¦i> commands revised. References to FLM-0177 board added. Reference to FLM-0177 board added. References to FLM-0177 board added. :SYSTem:PADJust commands revised. :SYSTem:PASSword:CENable:STATe? command revised. References to FLM-0177 board added. |
| Issue | Date | Section | Details |
|  |  | 5.1.11.2 5.1.12 5.1.12 5.1.13 5.1.14 5.1.14.1 5.1.15 5.1.16.1 5.1.16.1.1 5.1.17 5.1.17 5.1.17 5.1.17 5.1.19 5.1.19.2 5.1.20 5.2 5.3 5.4.2 5.4.3 5.5 5.5.1 Table 5.2 Table 5.3 Table 5.7 Table 5.11 Table 5.12 Table 5.14 Table 5.21 Table 5.24 Table 5.30 Table 5.31 Table 5.33 Table 5.36 | References to FLM-0177 board added. Reference to FLM-0177 board added. :SYSTem:PDRive:SFACtor commands added. References to FLM-0177 board added. References to FLM-0177 board added. Reference to FLM-0177 board added. Reference to FLM-0177 board added. References to FLM-0177 board added. Reference to FLM-0177 board added. Introduction to :SYSTem:TCONtrol<0[¦1¦2¦3¦4]> commands revised. References to FLM-0177 board added. :SYSTem:TCONtrol1:OMAX commands added. :SYSTem:TCONtrol1:OMIN commands added. References to FLM-0177 board added. Introduction to :SYSTem:VOA:MODulation commands revised. References to FLM-0177 board added. References to FLM-0177 board added. Reference to FLM-0177 board added. :SYSTem:DIAGnostic:TEST:EEPRom? command added. Reference to FLM-0177 board added Reference to FLM-0177 board added. Reference to FLM-0177 board added. Table updated. Table added. Table added. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. |
| AB | 20-Aug-15 | 5.1.1 5.1.13 5.1.13 5.1.13 5.1.13 Table 5.33 | :SYSTem:AMPLifier commands added. :SYSTem:PMONitor:P0 commands revised. :SYSTem:PMONitor:P1 commands revised. :SYSTem:PMONitor:POWer? command revised. :SYSTem:PMONitor:TFACtor commands revised. Table updated. |
| AA | 22-Nov-11 | 1.1 5.1.1 5.1.2 5.1.3.2 5.1.3.4 5.1.6 5.1.7.1 5.1.9 5.1.9 5.1.11 5.1.11 5.1.11.1 5.1.11.2 5.1.12 5.1.13 5.1.13 5.1.13 5.1.15 5.1.14 5.1.14.1 5.1.16.1 5.1.16.1.1 5.1.17 | Related documents updated. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. Reference to FLM-0151 board added. :SYSTem:PADJust:POWer:MUTE commands added. References to FLM-0151 board added. :SYSTem:PCONtrol:EMODulation commands added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. :SYSTem:PMONitor:POWer? command revised. :SYSTem:PMONitor:TFACtor commands revised. :SYSTem:RTIMe<0¦1> commands added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added. References to FLM-0151 board added |
| Issue | Date | Section | Details |
|  |  | 5.1.17 5.1.17 5.1.19 5.1.20 5.2 5.3 5.4.3 5.5 5.5.1 Table 5.11 Table 5.12 Table 5.21 Table 5.24 Table 5.26 Table 5.27 Table 5.30 Table 5.31 Table 5.33 Table 5.34 Table 5.37 | :SYSTem:TCONtrol<2>:DELay commands added. :SYSTem:TCONtrol<2>:VMONitor? command revised. :SYSTem:VOA commands added. References to FLM-0151 board added. :STATus commands revised. Reference to FLM-0151 board added. Reference to FLM-0151 board added. Reference to FLM-0151 board added. :OPTimization:LOG commands added. Table updated. Table updated. Table updated. Table updated. Table added. Table added. Table updated. Table updated. Table updated. Table updated. Table updated. |
| Z | 19-May-10 | 1.1 2 2.1 Figure 2.1 2.2 Figure 2.2 5.1.14.1 5.2 Table 5.33 Table 5.34 | Related documents updated. Section revised. Section revised. Figure revised. Section added. Figure added. :SYSTem:RIN:ADJust commands added. :STATus commands revised. Table updated. Table updated. |
| Y | 16-Mar-10 | 1.1 5.1.11 5.1.13 5.1.16.1 5.5.1 Table 5.20 Table 5.32 Table 5.33 Table 5.37 | Related documents updated. :SYSTem:PCONtrol:PLIMit commands added. :SYSTem:PMONitor:POWer? command revised. :SYSTem:SMODe:CORRection:STATe commands revised. :OPTimization:RAMP<0¦1>:SMRanges? command added. Table added. Table updated. Table updated. Table updated. |
| X | 29-Jun-09 | 1.1 2 3.2.9 5.1 5.1.1 5.1.3.1 5.1.3.3 5.1.6 5.1.7.1 5.1.7.1 5.1.7.1 5.1.7.1 5.1.8 5.1.11 5.1.11.1 5.1.11.1 5.1.11.2 5.1.12 5.1.13 5.1.13 5.1.14 5.1.14 5.1.14 5.1.16.1 | Related documents updated. Section updated. Section updated. :SYSTem:VREFerence<1¦2>? command added. :SYSTem:BFACet[0¦1] commands revised. :SYSTem:COMMunicate:ADDRess commands added. Introduction to:SYSTem:COMMunicate:SERial<0> commands revised. :SYSTem:HARDware:OPTions commands revised. :SYSTem:MHFTuning:BAND<i>:RTEMperature commands revised. :SYSTem:MHFTuning:BAND<i>:RVOLtage commands revised. :SYSTem:MHFTuning:BAND<i>:TSLope commands revised. :SYSTem:MHFTuning:BAND<i>:VSLope commands revised. Introduction to :SYSTem:OPTimization commands revised. :SYSTem:PCONtrol:CLIMit[0¦1] commands revised. :SYSTem:PCONtrol:MODE<0¦1>:RANGe commands revised. :SYSTem:PCONtrol:MODE<0¦1>:SPOint[0¦1] commands revised. :SYSTem:PCONtrol:RANGe<0[¦1¦2¦3]>:CMAX[0¦1] commands revised. Introduction to :SYSTem:PDRive commands revised. :SYSTem:PMONitor:C[0¦1]<0¦1> commands revised. :SYSTem:PMONitor:CURRent[0¦1]? command revised. :SYSTem:RIN:APASs commands added. :SYSTem:RIN:GAIN commands revised. :SYSTem:RIN:HPASs commands added. :SYSTem:SMODe:CORRection:HOLDoff commands revised |
| Issue | Date | Section | Details |
|  |  | 5.1.16.1 5.1.16.1.1 5.1.17 5.1.17 5.1.19 5.1.19 5.4.2 5.4.3 5.5 Table 5.12 Table 5.14 Table 5.16 Table 5.21 Table 5.24 Table 5.30 Table 5.31 Table 5.33 Table 5.36 | :SYSTem:SMODe:CORRection:STEP commands revised. Introduction to :SYSTem:SMODe:CORRection:RAMP commands revised. :SYSTem:TCONtrol<2>:DOUTput commands revised. :SYSTem:TCONtrol<2>:PWMPeriod commands added. :SYSTem:WADJust:INPut commands revised. :SYSTem:WADJust:STATe commands revised. :DIAGnostic:TEST:VREFerence commands added. :DIAGnostic:WDOG:BOOT command added. Introduction to :OPTimization commands revised. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. Table updated. |
| W | 16-Mar-09 | 1.1 5.1.8 5.4.3 5.5 Figure 5.3 Table 5.32 Table 5.33 Table 5.36 Table 5.37 | Related documents updated. :SYSTem:OPTimization commands added. :DIAGnostic:WDOG commands added. :OPTimization commands added. Figure added. Table added. Table updated. Table added. Table added. |
| V | 06-Oct-08 | 5.4.2 Table 5.37 | :DIAGnostic:TEST commands added. Table updated. |
| U | 04-Aug-08 | 1.1 3.2.9 Table 5.11 | Related documents updated. Section on multi-channel addressing added. Table updated. |
| T | 06-Feb-08 | Table 5.11 | Table updated. |
| S | 05-Dec-07 | 1.1 5.1.1 5.1.3.3 5.1.3.4 5.1.6 5.1.6 5.1.6 5.1.7 5.1.7.1 5.1.8 5.1.11 5.1.11 5.1.11 5.1.11.1 5.1.11.1 5.1.11.2 5.1.12 5.1.12.1 5.1.13 5.1.13 5.1.14 5.1.14.1 5.1.16.1 5.1.16.1 5.1.16.1 5.1.16.1 5.1.16.1.1 5.1.17 5.1.17 5.1.17 | Related documents updated. :SYSTem:BFACet[0¦1] commands revised. :SYSTem:COMMunicate:SERial<0> commands added. :SYSTem:COMMunicate:SERial<1> commands revised. :SYSTem:HARDware:LCDWelcome commands revised. :SYSTem:HARDware:OPTions commands revised. :SYSTem:HARDware:TYPE commands revised. :SYSTem:MHFTuning commands revised. :SYSTem:MHFTuning:BAND<a¦i> commands revised. :SYSTem:PADJust commands revised. :SYSTem:PCONtrol:CLIMit[0¦1] commands revised. :SYSTem:PCONtrol:DELay commands added. :SYSTem:PCONtrol:STATe commands added. :SYSTem:PCONtrol:MODE<0¦1>:RANGe commands revised. :SYSTem:PCONtrol:MODE<0¦1>:SPOint[0¦1] commands revised. :SYSTem:PCONtrol:RANGe<0[¦1¦2¦3]>:CMAX[0¦1] commands revised. :SYSTem:PDRive commands added. :SYSTem:PDRive:MODulation commands added. :SYSTem:PMONitor:C[0¦1]<0¦1> commands revised. :SYSTem:PMONitor:CURRent[0¦1]? command revised. :SYSTem:RIN:GAIN commands added. :SYSTem:SMODe:LEDDisable commands added. :SYSTem:SMODe:CORRection:FLAGs? command added. :SYSTem:SMODe:CORRection:HOLDoff commands added. :SYSTem:SMODe:CORRection:RANGe commands revised. :SYSTem:SMODe:CORRection:STEP commands revised. :SYSTem:SMODe:CORRection:RAMP commands added. :SYSTem:TCONtrol<2>:DOUTput commands added. :SYSTem:TCONtrol<2>:ELEMents<0¦1> commands added. :SYSTem:TCONtrol<2>:VMONitor<1¦2¦3>? command added. |
| Issue | Date | Section | Details |
|  |  | 5.2 5.3 5.4 Table 5.13 Table 5.16 Table 5.19 Table 5.21 Table 5.24 Table 5.30 Table 5.31 Table 5.33 Table 5.34 | :STATus commands added. :MESSage commands revised. :DIAGnostic commands added. Table updated. Table updated. Table added. Table updated. Table updated. Table added. Table added. Table updated. Table added. |
| R | 02-Aug-06 | 5.1.19 Table 5.28 Table 5.33 | :SYSTem:WADJust:FLAGs? command added. Table added. Table updated. |
| Q | 27-Jan-06 | 5.1.8 5.1.19 5.1.19 Table 5.33 Table 5.38 | :SYSTem:PADJust commands added. :SYSTem:WADJust commands revised. :SYSTem:WADJust:INPut commands added. Table updated. Errors +120 and +122 added. |
| P | 02-Nov-05 | 5.1.17 Table 5.21 Table 5.33 | :SYSTem:TCONtrol<0[¦1¦2¦3¦4] commands revised. Table updated. Table updated. |
| O | 07-Oct-05 | 5.1.16.1 Table 5.33 | :SYSTem:SMODe:CORRection:RANGe commands added. Table updated. |
| N | 27-Sep-05 | 5.1.6 Table 5.33 | :SYSTem:HARDware:LCDWelcome commands added. Table updated. |
| M | 22-Feb-05 | 1.1 5.1.3 5.3 Table 5.33 Table 5.35 | Related documents updated. :SYSTem:COMMunicate commands added. :MESSage commands added. Table updated. Table added. |
| L | 11-Oct-04 | 5.1.11.1 Table 5.33 | :SYSTem:PCONtrol:MODE<0¦1>:SPOint[0¦1] commands revised. Table updated. |
| K | 10-Sep-04 | 4.3 5.1.6 Table 5.11 | *OPT? command revised. :SYSTem:HARDware:OPTions commands revised. Table updated. |
| J | 25-Aug-04 | 5.1.6 5.1.16.1 Table 5.33 | :SYSTem:HARDware:APPLication commands added. :SYSTem:SMODe:CORRection commands added. Table updated. |
| I | 17-Aug-04 | 1.1 2 5.1.1 5.1.6 5.1.6 5.1.7.1 5.1.11 5.1.11.1 5.1.11.2 5.1.12 5.1.12 5.1.19 Table 5.12 Table 5.21 Table 5.23 Table 5.24 Table 5.33 | Related documents updated. Section on RS-232 operation added. :SYSTem:BFACet[0¦1] commands revised. :SYSTem:HARDware:OPTions <i> command revised. :SYSTem:HARDware:TYPE <i> command revised. :SYSTem:MHFTuning:BAND<a¦i> commands revised. :SYSTem:PCONtrol:CLIMit[0¦1] commands revised. :SYSTem:PCONtrol:MODE<0¦1>:RANGe commands revised. :SYSTem:PCONtrol:RANGe<0[¦1¦2¦3]>:CMAX[0¦1] commands revised. :SYSTem:PMONitor:C[0¦1]<0¦1> commands revised. :SYSTem:PMONitor:CURRent[0¦1]? command revised. :SYSTem:WADJust:STATe commands revised Table updated. Table updated. Table updated. Table updated. Table updated. |

| Issue | Date | Section | Details |
| --- | --- | --- | --- |
| H | 02-Jun-04 | 1 1.1 5.1.1 5.1.11 5.1.11.1 5.1.11.2 5.1.12 5.1.14 5.1.14.1 5.1.17 5.1.18 5.1.19 Table 5.12 Table 5.14 Table 5.21 Table 5.24 Table 5.33 | Introduction updated. Related documents updated. :SYSTem:BFACet commands added. :SYSTem:PCONtrol:CLIMit commands added. :SYSTem:PCONtrol:MODE<0¦1>:RANGe commands revised. :SYSTem:PCONtrol:RANGe<0[¦1¦2¦3]> commands revised. :SYSTem:PMONitor commands revised. Section updated. Section updated. :SYSTem:TCONtrol<0[¦1¦2¦3]> commands revised. :SYSTem:TMONitor<0[¦1¦2¦3¦4]> commands revised. Section updated. Table added. Table updated. Table updated. Table updated. Table updated. |
| G | 14-Apr-04 | 1.1 5.1.19 5.1.19 Table 5.33 Table 5.38 | Related documents updated. :SYSTem:WADJust section revised. :SYSTem:WADJust:WAVelength commands added. Table updated. Errors +110, +111, +112 and +113 added. |
| F | 22-Mar-04 | 5.1.6 5.1.6 5.1.6 5.1.12 5.1.12 5.1.16.1 5.1.18 Table 5.21 Table 5.23 Table 5.24 Table 5.33 | :SYSTem:HARDware:OPTions commands added. :SYSTem:HARDware:REVision commands added. :SYSTem:HARDware:TYPE commands added. :SYSTem:PMONitor:C[0¦1]<0¦1> commands revised. :SYSTem:PMONitor:CURRent[0¦1]? command revised. :SYSTem:TCONtrol<0¦1¦2[¦3]> commands revised. :SYSTem:TMONitor<0¦1¦2¦3[¦4]> commands revised. Table updated. Table updated. Table updated. Table updated. |
| E | 05-Mar-04 | 3.2.1 5.1.19 Table 5.33 | Alpha data type added. :SYSTem:WADJust commands added. Table updated. |
| D | 26-Feb-04 | 5.1.11 5.1.11.1 5.1.11.1 5.1.11.1 5.1.11.1 Table 5.33 | :SYSTem:PCONtrol:SPOint commands revised. :SYSTem:PCONtrol:MODE<0>:KD commands added. :SYSTem:PCONtrol:MODE<0>:KI commands added. :SYSTem:PCONtrol:MODE<0>:KP commands added. :SYSTem:PCONtrol:MODE<0>:SPOint commands revised. Table updated. |
| C | 11-Feb-04 | 5.1 5.1.12 Table 5.33 | :SYSTem:ADCOffset commands added. :SYSTem:PMONitor:TFACtor commands added. Table updated. |
| B | 21-Jan-04 | 5.1.7 Table 5.33 | :SYSTem:MHFTuning commands added. Table updated. |
| A | 07-Jan-04 | n/a | First issue. |

# Introduction

This document describes the command language used to set up and control the FLM-0100 series fiber laser controllers from a host computer. It includes an overview of the language, and detailed descriptions of each of the commands.

## Related Documents

  - Fiber Laser Controller – Hardware Design Report	FLM-003, Rev. D

  - Fiber Laser Controller – Test Utility User’s Guide	FLM-006, Rev. A

  - Twin Pump Laser – Hardware Design Report	FLM-009, Rev. B

  - ASE Controller – Hardware Design Report	FLM-010, Rev. A

  - Fiber Laser Controller – FLM-0106 Hardware Design Report	FLM-011, Rev. B

  - Fiber Laser Controller – Control Panel User’s Guide	FLM-012, Rev. A

  - Fiber Laser Controller – FLM-0107 Hardware Design Report	FLM-014, Rev. B

  - Fiber Laser Controller – FLM-0121 / FLM-0125 Design Report	FLM-042, Rev. B

  - Fiber Laser Controller – FLM-0130 Hardware Design Report	FLM-044, Rev. F

  - Fiber Laser Controller – Watchdog Diagnostics Commands	FLM-067, Rev. A

  - Fiber Laser Controller – Optimization Facilities	FLM-073, Rev. B

  - Fiber Laser Controller – FLM-0141 Hardware Design Report	FLM-075, Rev. F

  - Fiber Laser Controller – FLM-0151 Hardware Design Report	FLM-093, Rev. E

  - Fiber Laser Controller – FLM-0177 Hardware Design Report	FLM-0177R, Rev. A

  - FLM-0178 20A Diode Driver – Hardware Design Report	FLM-0178R, Rev. A

  - Power Rock System – Hardware Design Report	FLM-0181R, Rev. L

  - FLM-0194 5W Gen V System – Hardware Design Report	FLM-0194R, Rev. D

  - FLM-0197 Amplified ASE – Hardware Design Report	FLM-0197R, Rev. A

  - Power Rock – Firmware & Software Settings	FLM-190, Rev. E

  - Power Rock – User Command Language	FLM-192, Rev. C

  - Amplified ASE – Firmware & Software Settings	FLM-199, Rev. A

  - 690039 (Adressing in NP Multi-Board Systems)	690039 Rev 1

  - PR600-686 (Procedure, GEN VI Setup)	PR600686 Rev 1

  - 600637 (GEN VI HPLS,FLCP GUI REQUIREMENTS)	600637 Rev 4

## References

  - Standard Commands for Programmable Instrumentation (SCPI)	http://www.scpiconsortium.org

# Hardware Interface

The fiber laser controller communicates via its RS-232 serial port, which can be connected to a host PC, either directly or indirectly. Indirect connection options include USB and Ethernet converters, which generally use virtual COM port drivers to establish a conventional serial link. Where appropriate, the port on the PC should be set up for 8 data bits, 1 stop bit, no parity and no flow control. The baud rate should be set to 57600. The transmit and receive signals at the laser and the interface module are at RS-232 voltage levels.

## RS-232 Connections – via Interface Module

For connections to a host PC, using the interface module supplied with the laser, an RS-232 extension cable should be used. This is a DB-9 male to female cable assembly with one-to-one pin connections. There should be no crossover of the transmit and receive signals. Note that a null modem cable is not suitable.

Figure 2.1 – RS-232 Connector Pin-Out (Interface Module)

|  | Pin; Function. 1; no connection. 2; TXD, transmit data (output from laser). 3; RXD, receive data (input to laser). 4; no connection. 5; GND, signal ground. 6; no connection. 7; no connection. 8; no connection. 9; no connection. |
| --- | --- |

The connector pin-out is illustrated in Figure 2.1. Note that the serial port uses the transmit (TXD), receive (RXD) and signal ground (GND) lines only; the pins normally used for hardware handshaking (Pin 7 - CTS and Pin 8 - RTS) are not connected.

## RS-232 Connections – direct to Laser Module

If the standard interface module is not being used, then it is necessary to provide certain signals and voltages to the DB-25 female connector on the laser module.

Figure 2.2 – RS-232 Connector Pin-Out (Laser Module)

|  | Pin; Function. 1; GND, ground connection. 2; GND, ground connection. 3; GND, ground connection. 14; GND, ground connection. 15; GND, ground connection. 24; TXD, transmit data (output from laser). 25; RXD, receive data (input to laser). others; see manual. |
| --- | --- |

The connector pin-out is illustrated in Figure 2.2. Note that five pins are used for power and signal grounds on the laser interface connector. A connection should be made from the host PC’s RS-232 GND to any one of these pins.

Note too that the above table only shows the pins associated with the RS-232 communications. For correct operation of the laser, it is also necessary to make connections to a number of other pins. Refer to the manual supplied with the laser for details.

# Command Language Overview

## General

The command language is based on the Standard Commands for Programmable Instrumention (SCPI) specification y, and offers the following advantages:

- Command messages follow a standard, structured format, similar to those used for other programmable instruments.

- Command messages are comprised of command words (or mnemonics) and parameters, in ASCII text format, allowing easy control from most programming environments, including LabView and Visual Basic, or even using a terminal emulator such as Windows HyperTerminal.

- New commands can be added to the language relatively easily to support on-going development.

Note:	Although the command language is based on the SCPI style, it does not support all of the features defined in the standard, and no claim is made to the effect that it is SCPI compliant.

## Command Syntax

### Commands and Command Parameters

Commands may or may not use a parameter. Where a parameter is used, it should be separated from the command word by at least one space. Unless explicitly noted, parameter values are specified in ASCII text format. A number of different data types are used, as follows:

<a>	An alpha character, i.e. a single letter. This can be either upper or lower case.

<b>	A Boolean value. Used to enable or disable the specified function. 1 enables the function, and 0 disables it. The values ON, OFF, ENABled, DISabled, may also be used, if preferred.

<f>	A floating-point numeric value. These parameters may be expressed using either decimal format (e.g. 0.123) or exponential format (e.g. 1.23E-1).

<i>	An integer value (e.g. 123).

<s>	A character string. Character strings are enclosed within double quotes (“…”).

Note that the angle brackets (<>) are used to denote the type of parameter and should not be included in the command message. For example, in the message

:SYSTem:TCONtrol0:SPOint <f>

<f> is used to indicate that a floating-point parameter is required. Thus, to set the Temperature Control 0 set-point to 25ºC, the message would be sent as:

:SYSTem:TCONtrol0:SPOint 25.0

### Query Commands

This type of command is used to request parameter values or data from the controller. It is identified by a question mark (?) immediately following the command word. All of the commands that allow parameter values to be specified also have a query form, providing the means to report the value currently set.

### Case Sensitivity

All commands are non-case-sensitive. Messages can be sent using either upper case or lower case, or a combination of both.

### Long-Form and Short-Form

Many of the SCPI-style command words have both long-form and short-form versions. For example, the following messages are all valid:

:SYSTEM:FIRMWARE:VERSION?		long-form

:SYST:FIRM:VERS?					short-form

:SYST:FIRM:VERSION?				combination of short-form and long-form

The Command Reference in Section 5 defines the long-form version of each command, but uses upper case characters to indicate the short-form version. Thus, the message illustrated above is defined in the form :SYSTem:FIRMware:VERSion?

Note that each command word must be in either its long-form or short-form and not something in between. For example, the message :SYST:FIRMw:VERSio? is invalid and will generate an error.

### Program Messages

A program message consists of one or more command words and associated parameters, sent by the host computer to the controller. Command words are separated by a colon (:), giving a multi-layer structure to the message. This is best illustrated by example:

:SYSTem						path (root)

:PCONtrol				path

:MODE1					path

:SPOint <i>				command word and parameter

:SPOint?					query command

In this example there are four levels. The first is the root command and serves as a path to the SYSTEM command group. The second is also a path, this time to the group of commands associated with the Pump Control. The third is another path, to the group of commands associated with operating Mode 1 (constant current). In the fourth level, there are two commands, one for specifying the set-point (in bits) and the other for querying it. The two commands in this example can be executed by issuing the following two program messages:

:SYST:PCON:MODE1:SPO <i>

:SYST:PCON:MODE1:SPO?

Note:	The SCPI standard also provides for sending multiple command messages in the same program message, by separating the individual command messages by semicolons (;). However, the fiber laser controllers do not currently support this feature.

### Program Message Terminator

Every program message must be terminated by a carriage return - line feed sequence (<CR><LF>).

### Response Messages

A response message is the message sent by the controller to the host computer in response to a program message containing a query command. Parameter and data values are returned in ASCII text format, unless noted otherwise. Data types are as used for the command parameters.

### Response Message Terminator

Every response message is terminated by a carriage return - line feed sequence (<CR><LF>).

### Multi-Channel Addressing

In multi-channel systems, commands are routed to all channels simultaneously, via a multi-drop bus. Each controller is assigned a unique address and will only respond to commands sent to that address. The channel address is sent as a prefix to the command message, using one of the formats described below. Response messages are exactly as for the single channel systems. For details on configuring the address settings, refer to Section 5.1.3.1.

#### Physical Addressing

The FLM-0121, FLM-0125 and FLM-0141 boards each have a physical address defined by the hardware and may be addressed using the format,

#<physical address>,<command>

where <physical address> can be in the range 0 to 9 (FLM-0121 and FLM-0125) or 0 to 127 (FLM-0141) and <command> can be any SCPI program message or common command.

#### Aliased Addressing

All board types can be assigned an address alias, which can be used to provide a meaningful name or reference for each channel. On the FLM-0121, FLM-0125 and FLM-0141 boards, the alias can then be used instead of the hardware address. On the other board types, the alias makes it possible to add an addressing capability to individual modules. In all cases, boards are addressed using the format,

@<alias address>,<command>

where <alias address> is the alias and <command> can be any SCPI program message or common command.

# Common Commands

The common commands are a small group of commands that use a different format to the SCPI commands and perform basic functions such as unit identification, reset and status operations. They are normally associated with instruments that can be controlled via a GPIB interface, and are defined as part of the IEEE 488.2 standard. However, instruments provided with both GPIB and serial interfaces usually support the common commands via both interfaces. Therefore, for consistency, the relevant commands are implemented on the FLM-0100 series controllers in a similar manner. These are listed in Table 4.1 and described below.

Table 4.1 – Common Command Summary

| Command | Description |
| --- | --- |
| *CLS | Clear status (clears any errors) |
| *IDN? | Query unit identification |
| *OPT? | Query option identification |
| *RST | Reset to power-on state (software reset) |
| *SAV | Save configuration to non-volatile memory |

## *CLS – Clear Status

This command is used to clear the status of the controller. Any errors will be cleared (the Error LED will be extinguished) and the controller will be left ready to accept another command.

## *IDN? – Identification Query

This command returns the controller’s identification string, which comprises the manufacturer, model number, serial number and firmware version number. The string will be in the form:

NP PHOTONICS,FLM-0100,ssssssss,vvvvvvv

where	ssssssss is the serial number and vvvvvvv is the firmware version number.

## *OPT? – Option Identification Query

This command can be used to determine whether any hardware or firmware options are present. If the high power option is fitted, “HP” is reported; if there are no options fitted, “0” is reported.

## *RST – Reset

This command performs a software reset, returning the controller to its power-up state.

## *SAV – Save

This command saves the controller’s present configuration into non-volatile memory. It is intended for use when setting up the unit in manufacture. Refer to Section 5.1.10 for further details.

# SCPI Command Reference

The command set is organized into a number of groups, which are summarized in the tables listed below and described in detail in the remainder of this document.

SYSTem command summary				Table 5.33

STATus command summary				Table 5.34

MESSage command summary				Table 5.35

DIAGnostic command summary				Table 5.36

OPTimization command summary			Table 5.37

The following general comments apply:

- Upper case characters are used to indicate the short form version of each command word. Refer to Section 3.2.4 for further details.

- Square brackets ([]) are used to indicate optional characters, which do not have to be present in the program message. The brackets should not be included in the program message.

- Angle brackets (<>) are used to indicate the parameter type required. The brackets should not be included in the program message. Refer to Section 3.2.5 for further details.

- A vertical bar (|) is used to separate two or more alternatives in a list of possible values.

- A checkmark () in the Config. column indicates that the parameter forms part of the controller’s configured data set and, as such, can be saved into non-volatile memory and restored each time the unit boots up. Refer to Sections 4.5 and 5.1.8 for further details.

- The Version column is used to indicate the earliest version of firmware required to support the command. Unless noted otherwise, backwards compatibility is maintained.

## :SYSTem Commands

These commands are used to set up and query the system settings. They include commands for setting up and reading back the controller’s control settings and hardware/firmware details, password-protecting the configured settings, and reporting errors.

###### :ADCOffset <i>

This command is used to specify the ADC offset. It is intended for factory test purposes only and allows any offset in the ADC input circuitry to be trimmed out.

Syntax:	:SYSTem:ADCOffset <i>

Parameters:	<i>	ADC offset in bits

Configured:	yes

Response:	none

###### :ADCOffset?

This command is used to query the ADC offset.

Syntax:	:SYSTem:ADCOffset?

Parameters:	none

Response:	returns the ADC offset in bits

###### :VREFerence<1|2>?

This command is used to query the 1.042 V and 2.048 V reference signals measured by the ADC, allowing the ADC scale factor and offset to be calibrated. It is available on the FLM-0141 board only.

Syntax:	:SYSTem:VREFerence<1|2>?

Parameters:	none

Response:	returns the reference voltage in bits

### :AMPLifier Commands

This group of commands is used to set up and query the amplifier settings. Note that they are only available on the Gen IV (hardware application = “HPS-4”) and Power Rock (hardware application = “HPS-5” and “HPS-5A”) high power systems and the Amplified ASE (hardware application = HPASE).

###### :C[0|1|2]0 <f>

###### :C[0|1|2]1 <f>

These commands are used to specify current coefficients C0 and C1 on the Gen IV system, current coefficients C00, C01, C10, C11, C20 and C21 on the Power Rock systems, or current coefficients C00, C01, C10, and C11 on the Amplified ASE.

Syntax:	:SYSTem:AMPLifier:C0 <f>	Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:C1 <f>	Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:C<0|1>0 <f>	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:C<0|1>1 <f>	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:C<0|1|2>0 <f>	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:C<0|1|2>1 <f>	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:C<0|1>0 <f>	Amp. ASE	(f/w version 8.5.000 onwards)

:SYSTem:AMPLifier:C<0|1>1 <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	current drive coefficient C[0]0 (bits), or C[0]1 (bits/A), as appropriate

current monitor coefficient C10 (A), or C11 (A/bit), as appropriate

Stage 2 current drive coefficient C20 (bits), or C21 (bits/A)	(2-stage Power Rock)

Configured:	yes

Response:	none

###### :C[0|1|2]0?

###### :C[0|1|2]1?

These commands are used to query current coefficients C0 and C1 on the Gen IV system, current coefficients C00, C01, C10, C11, C20 and C21 on the Power Rock systems, or current coefficients C00, C01, C10, and C11 on the Amplified ASE.

Syntax:	:SYSTem:AMPLifier:C0?		Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:C1?		Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:C<0|1>0?	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:C<0|1>1?	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:C<0|1|2>0?	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:C<0|1|2>1?	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:C<0|1>0?	Amp. ASE	(f/w version 8.5.000 onwards)

:SYSTem:AMPLifier:C<0|1>1?	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns	current drive coefficient C[0]0 (bits), or C[0]1 (bits/A), as appropriate

or	current monitor coefficient C10 (A), or C11 (A/bit), as appropriate

or	Stage 2 current drive coefficient C20 (bits), or C21 (bits/A)	(2-stage Power Rock)

###### :CMAXimum<[0]|1> <f>

This command is used to specify the maximum current set-point(s).

Syntax:	:SYSTem:AMPLifier:CMAXimum <f>		(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:CMAXimum<[0]|1> <f>	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:CMAXimum <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	maximum current set-point in A

maximum Stage 2 current set-point (CMAXimum[0]) in A	(2-stage Power Rock)

maximum Stage 1 current set-point (CMAXimum1) in A	(2-stage Power Rock)

Configured:	yes

Response:	none

###### :CMAXimum<[0]|1>?

This command is used to query the maximum current set-point(s).

Syntax:	:SYSTem:AMPLifier:CMAXimum?		(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:CMAXimum<[0]|1>?	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:CMAXimum?	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns	the maximum current set-point in A

or	the maximum Stage 2 current set-point (CMAXimum[0]) in A	(2-stage Power Rock)

or	the maximum Stage 1 current set-point (CMAXimum1) in A	(2-stage Power Rock)

###### :CMINimum<[0]|1> <f>

This command is used to specify the minimum current set-point(s).

Syntax:	:SYSTem:AMPLifier:CMINimum <f>		(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:CMINimum<[0]|1> <f>	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:CMINimum <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	minimum current set-point in A

minimum Stage 2 current set-point (CMINimum[0]) in A	(2-stage Power Rock)

minimum Stage 1 current set-point (CMINimum1) in A	(2-stage Power Rock)

Configured:	yes

Response:	none

###### :CMINimum<[0]|1>?

This command is used to query the minimum current set-point(s).

Syntax:	:SYSTem:AMPLifier:CMINimum?		(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:CMINimum<[0]|1>?	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:CMINimum?	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns	the minimum current set-point in A

or	the minimum Stage 2 current set-point (CMINimum[0]) in A	(2-stage Power Rock)

or	the minimum Stage 1 current set-point (CMINimum1) in A	(2-stage Power Rock)

###### :CMONitor?

This command is used to query the monitored current. On the 2-stage Power Rock, it refers to the first stage (Stage 1).

Syntax:	:SYSTem:AMPLifier:CMONitor?

Parameters:	none

Response:	returns the monitored current in A

###### CSETpoint<[0]|1> <f>

This command is used to specify the current set-point(s).

Syntax:	:SYSTem:AMPLifier:CSETpoint <f>		(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:CSETpoint<[0]|1> <f>	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:CSETpoint <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	current set-point in A

Stage 2 current set-point (CSETpoint[0]) in A	(2-stage Power Rock)

Stage 1 current set-point (CSETpoint1) in A	(2-stage Power Rock)

Configured:	yes

Response:	none

###### :CSETpoint<[0]|1>?

This command is used to query the current set-point(s).

Syntax:	:SYSTem:AMPLifier:CSETpoint?		(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:CSETpoint<[0]|1> <f>	Power Rock	(f/w version 8.4.000 onwards)

:SYSTem:AMPLifier:CSETpoint?	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns	the current set-point in A

or	the Stage 2 current set-point (CSETpoint[0]) in A	(2-stage Power Rock)

or	the Stage 1 current set-point (CSETpoint1) in A	(2-stage Power Rock)

###### :FLAGs?

This command is used to query the amplifier control state. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:AMPLifier:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.2

Table 5.2 – Amplifier Status Flags

| Bit | Description |
| --- | --- |
| 0-2 | Operating Mode (0 = APC; 1 = ACC) |
| 3 | Disabled |
| 4 | Output active |
| 5 | Shutdown |
| 6 | reserved |
| 7 | reserved |
| 8 | H/W enable |
| 9 | S/W enable |
| 10 | Low input power |
| 11 | High SBS power |
| 12-15 | Shutdown code (see Table 5.3) |

Note:	All bits are active high.

Shutdown code requires f/w version 8.1.000 onwards.

Table 5.3 – Amplifier Shutdown Codes

| Code | Description |
| --- | --- |
| 0 | Normal (no shutdown) |
| 1 | Amplifier Disabled (H/W or S/W) |
| 2 | Seed Disabled (H/W or S/W) |
| 3 | Seed Tuning |
| 4 | Seed Optimizing |
| 5 | Seed Temperature Fault |
| 6 | Amplifier Temperature Fault |
| 7 | reserved |
| 8 | Emergency Stop / Remote Interlock |
| 9 | Low Input Power |
| 10 | High SBS Power |
| 11-15 | reserved |

Note:	Codes 0 - 7 are slow shutdown conditions.

Codes 8 - 15 are fast shutdown conditions.

Code 10 not used on 2-stage Power Rock.

###### :MSELect <i>

This command is used to specify the operating mode.

Syntax:	:SYSTem:AMPLifier:MSELect <i>

Parameters:	<i>	operating mode	see Table 5.4

Configured:	yes

Response:	none

Table 5.4 – Amplifier Operating Modes

| Mode | Description |
| --- | --- |
| 0 | Constant power (APC) |
| 1 | Constant current (ACC) |

###### :MSELect?

This command is used to query the operating mode.

Syntax:	:SYSTem:AMPLifier:MSELect?

Parameters:	none

Response:	returns the operating mode	see Table 5.4

###### :OPTions <i>

This command is used to specify the amplifier options. The parameter value is an integer comprising a series of flags.

Syntax:	:SYSTem:AMPLifier:OPTions <i>

Parameters:	<i>	amplifier options	see Table 5.4

Configured:	yes

Response:	none

Table 5.5 – Amplifier Options

| Bit | Description |
| --- | --- |
| 0 | No APC |
| 1 | Shutdown while seed is tuning |
| 2 | Shutdown while seed is optimizing |
| 3-15 | reserved |

Note:	All bits are active high.

###### :OPTions?

This command is used to query the amplifier options.

Syntax:	:SYSTem:AMPLifier:OPTions?

Parameters:	none

Response:	returns the amplifier options	see Table 5.4

###### :P[0|1]0 <f>

###### :P[0|1]1 <f>

These commands are used to specify power monitor coefficients P0 and P1 on the Gen IV system, or power monitor coefficients P00, P01, P10 and P11 on the Power Rock systems. On the 2-stage Power Rock, the output power monitor refers to Stage 2 and the SBS power monitor is used for Stage 1.

Syntax:	:SYSTem:AMPLifier:P0 <f>	Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:P1 <f>	Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:P<0|1>0 <f>	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:P<0|1>1 <f>	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:P<0>0 <f>	Amp. ASE	(f/w version 8.5.000 onwards)

:SYSTem:AMPLifier:P<0>1 <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	output power monitor coefficient P[0]0 (mW), or P[0]1 (mW/bit), as appropriate

SBS/Stage 1 power monitor coefficient P10 (mW), or P11 (mW/bit), as appropriate

Configured:	yes

Response:	none

###### :P[0|1]0?

###### :P[0|1]1?

These commands are used to query power monitor coefficients P0 and P1 on the Gen IV system, or power monitor coefficients P00, P01, P10 and P11 on the Power Rock systems. On the 2-stage Power Rock, the output power monitor refers to Stage 2 and the SBS power monitor is used for Stage 1.

Syntax:	:SYSTem:AMPLifier:P0? 		Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:P1? 		Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:P<0|1>0?	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:P<0|1>1?	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:P<0>0?	Amp. ASE	(f/w version 8.5.000 onwards)

:SYSTem:AMPLifier:P<0>1?	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns	output power monitor coefficient P[0]0 (mW), or P[0]1 (mW/bit), as appropriate

or	SBS/Stage 1 power monitor coefficient P10 (mW), or P11 (mW/bit), as appropriate

###### :PMAXimum <f>

This command is used to specify the maximum power set-point. On the 2-stage Power Rock, it refers to the second stage (Stage 2).

Syntax:	:SYSTem:AMPLifier:PMAXimum <f>

Parameters:	<f>	maximum power set-point in W

Configured:	yes

Response:	none

###### :PMAXimum?

This command is used to query the maximum power set-point. On the 2-stage Power Rock, it refers to the second stage (Stage 2).

Syntax:	:SYSTem:AMPLifier:PMAXimum?

Parameters:	none

Response:	returns the maximum power set-point in W

###### :PMINimum <f>

This command is used to specify the minimum power set-point. On the 2-stage Power Rock, it refers to the second stage (Stage 2).

Syntax:	:SYSTem:AMPLifier:PMINimum <f>

Parameters:	<f>	minimum power set-point in W

Configured:	yes

Response:	none

###### :PMINimum?

This command is used to query the minimum power set-point. On the 2-stage Power Rock, it refers to the second stage (Stage 2).

Syntax:	:SYSTem:AMPLifier:PMINimum?

Parameters:	none

Response:	returns the minimum power set-point in W

###### :PMONitor<[0]|1>?

This command is used to query the monitored power levels. On the 2-stage Power Rock, the output power monitor refers to Stage 2 and the SBS power monitor is used for Stage 1.

Syntax:	:SYSTem:AMPLifier:PMONitor<[0]|1>?

Parameters:	none

Response:	returns the monitored power	see Table 5.6

Table 5.6 – Amplifier Power Monitors

| Monitor | Description | Units |
| --- | --- | --- |
| 0 | Output power monitor1 | W |
| 1 | SBS/Stage 1 power monitor2 | see Note 3 |

1 Stage 2 output on 2-stage Power Rock

2 Stage 1 output on 2-stage Power Rock

3 Units are mW (SBS) or W (Stage 1)

###### :PSETpoint <f>

This command is used to specify the power set-point. On the 2-stage Power Rock, it refers to the second stage (Stage 2).

Syntax:	:SYSTem:AMPLifier:PSETpoint <f>

Parameters:	<f>	power set-point in W

Configured:	yes

Response:	none

###### :PSETpoint?

This command is used to query the power set-point. On the 2-stage Power Rock, it refers to the second stage (Stage 2).

Syntax:	:SYSTem:AMPLifier:PSETpoint?

Parameters:	none

Response:	returns the power set-point in W

###### :RRATe<[0]|1> <f>

This command is used to specify the current set-point ramp rate(s).

Syntax:	:SYSTem:AMPLifier:RRATe <f>		(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:RRATe<[0]|1> <f>	Power Rock	(f/w version 8.4.002 onwards)

:SYSTem:AMPLifier:RRATe <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	current set-point ramp rate in A/s/10

Stage 2 current set-point ramp rate (RRATe[0]) in A/s/10	(2-stage Power Rock)

Stage 1 current set-point ramp rate (RRATe1) in A/s/10	(2-stage Power Rock)

Configured:	yes

Response:	none

###### :RRATe<[0]|1>?

This command is used to query the current set-point ramp rate(s).

Syntax:	:SYSTem:AMPLifier:RRATe?		(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:RRATe<[0]|1>?	Power Rock	(f/w version 8.4.002 onwards)

:SYSTem:AMPLifier:RRATe?	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns	the current set-point ramp rate in A/s/10

or	the Stage 2 current set-point ramp rate (RRATe[0]) in A/s/10	(2-stage Power Rock)

or	the Stage 1 current set-point ramp rate (RRATe1) in A/s/10	(2-stage Power Rock)

###### :STATe <b>

This command is used to specify the software enable state.

Syntax:	:SYSTem:AMPLifier:STATe <b>

Parameters:	<b>	s/w enable state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the software enable state.

Syntax:	:SYSTem:AMPLifier:STATe?

Parameters:	none

Response:	returns the s/w enable state	0 = disabled; 1 = enabled

###### :TFACtor[0|1] <f>

This command is used to specify the tap factors associated with the output power monitor (Gen IV and Power Rock) and SBS power monitor (Power Rock only). The tap factor is the reciprocal of the tap ratio. For example, a 0.1% (30dB) tap would require a tap factor of 1000.

Syntax:	:SYSTem:AMPLifier:TFACtor <f>	Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:TFACtor<0|1> <f>	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:TFACtor<0> <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	tap factor 	see Table 5.7

Configured:	yes

Response:	none

Table 5.7 – Amplifier Tap Factors

| Factor | Description |
| --- | --- |
| 0 | Output power monitor |
| 1 | SBS power monitor |

###### :TFACtor[0|1]?

This command is used to query the tap factors associated with the output power monitor (Gen IV and Power Rock) and SBS power monitor (Power Rock only).

Syntax:	:SYSTem:AMPLifier:TFACtor?	Gen IV	(f/w version 7.5.000 onwards)

:SYSTem:AMPLifier:TFACtor<0|1>?	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:AMPLifier:TFACtor<0>?	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns the tap factor 	see Table 5.7

###### :THReshold<0|1> <f>

This command is used to specify the power thresholds. It does not apply to the Amplified ASE.

Syntax:	:SYSTem:AMPLifier:THReshold<0|1> <f>

Parameters:	<f>	power threshold	see Table 5.8

Configured:	yes

Response:	none

Table 5.8 – Amplifier Power Thresholds

| Threshold | Description | Units |
| --- | --- | --- |
| 0 | Minimum input power | mW |
| 1 | Maximum SBS power1 | mW |

1 Threshold not available on 2-stage Power Rock

###### :THReshold<0|1>?

This command is used to query the power thresholds. It does not apply to the Amplified ASE.

Syntax:	:SYSTem:AMPLifier:THReshold<0|1>?

Parameters:	none

Response:	returns the power threshold 	see Table 5.8

###### :TMONitor<0|1>?

This command is used to query the monitored temperatures.

Syntax:	:SYSTem:AMPLifier:TMONitor<0|1>?

Parameters:	none

Response:	returns the temperature	see Table 5.9

Table 5.9 – Amplifier Temperature Monitors

| Monitor | Description | Units |
| --- | --- | --- |
| 0 | PCB temperature | °C |
| 1 | TEC temperature | °C |

#### :CONTrol Commands

This group of commands is used to set up and query the power control settings.

###### :KD <f>

This command is used to specify the differential constant used by the power control loop.

Syntax:	:SYSTem:AMPLifier:CONTrol:KD <f>

Parameters:	<f>	differential constant in bits/W.s

Configured:	yes

Response:	none

###### :KD?

This command is used to query the differential constant used by the power control loop.

Syntax:	:SYSTem:AMPLifier:CONTrol:KD?

Parameters:	none

Response:	returns the differential constant in bits/W.s

###### :KI <f>

This command is used to specify the integral constant used by the power control loop.

Syntax:	:SYSTem:AMPLifier:CONTrol:KI <f>

Parameters:	<f>	integral constant in bits/W/s

Configured:	yes

Response:	none

###### :KI?

This command is used to query the integral constant used by the power control loop.

Syntax:	:SYSTem:AMPLifier:CONTrol:KI?

Parameters:	none

Response:	returns the integral constant in bits/W/s

###### :KP <f>

This command is used to specify the proportional constant used by the power control loop.

Syntax:	:SYSTem:AMPLifier:CONTrol:KP <f>

Parameters:	<f>	proportional constant in bits/W

Configured:	yes

Response:	none

###### :KP?

This command is used to query the proportional constant used by the power control loop.

Syntax:	:SYSTem:AMPLifier:CONTrol:KP?

Parameters:	none

Response:	returns the proportional constant in bits/W

### :BFACet[0|1] Commands

This group of commands is used to set up and query the back facet settings. They are available on the FLM-0105, FLM-0106, FLM-0107, FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards. On the twin pump boards (FLM-0107, FLM-130, FLM-141, FLM-0151 and FLM-0177), separate back facet settings are provided for each pump.

###### :FLAGs?

This command is used to query the back facet signal status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:BFACet:FLAGs?		FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:BFACet:FLAGs?		FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:BFACet<0|1>:FLAGs?	FLM-0107	(f/w version 4.1.000 onwards)

:SYSTem:BFACet:FLAGs?		FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:FLAGs?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:FLAGs?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:BFACet<0|1>:FLAGs?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:BFACet<0|1>:FLAGs?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the status flags	see Table 5.10

Table 5.10 – Back Facet Flags

| Bit | Description |
| --- | --- |
| 0 | Back facet signal detected |

Note:	All bits are active high.

###### :LEVel?

This command is used to query the back facet signal level.

Syntax:	:SYSTem:BFACet:LEVel?		FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:BFACet:LEVel?		FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:BFACet<0|1>:LEVel?	FLM-0107	(f/w version 4.1.000 onwards)

:SYSTem:BFACet:LEVel?		FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:LEVel?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:LEVel?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:BFACet<0|1>:LEVel?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:BFACet<0|1>:LEVel?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the back facet signal level in bits

###### :THReshold <i>

This command is used to specify the back facet signal threshold. When the back facet signal level is at or above this value, the associated flag will be set.

Syntax:	:SYSTem:BFACet:THReshold <i>	FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:BFACet:THReshold <i>	FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold <i>	FLM-0107	(f/w version 4.1.000 onwards)

:SYSTem:BFACet:THReshold?	FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	<i>	back facet signal threshold in bits

Configured:	yes

Response:	none

###### :THReshold?

This command is used to query the back facet signal threshold.

Syntax:	:SYSTem:BFACet:THReshold?	FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:BFACet:THReshold?	FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0107	(f/w version 4.1.000 onwards)

:SYSTem:BFACet:THReshold?	FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:BFACet<0|1>:THReshold?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the back facet signal threshold in bits

### :COMMunicate Commands

This group of commands is used to set up and query the data communications settings.

#### :ADDRess Commands

This group of commands is used to set up and query the address settings. Addressing applies principally to the FLM-0121, FLM-0125 and FLM-0141 boards, which are designed for multi-channel applications. These boards have unique hardware addresses, so that commands can be directed to a particular channel. It is also possible to assign an alias to each one, allowing more meaningful address identifiers to be used. Address aliases can also be assigned to the other controller types, providing the means to address them in a similar manner. This makes it possible to construct a multi-channel system using the older-style modules. For details on address syntax, refer to Section 3.2.9.

###### :ALIas <s>

This command is used to specify an address alias.

Syntax:	:SYSTem:COMMunicate:ADDRess:ALIas <s>

Parameter:	<s>	alias as a string of up to 15 characters

Configured:	yes

Response:	none

###### :ALIas?

This command is used to query the address alias.

Syntax:	:SYSTem:COMMunicate:ADDRess:ALIas?

Parameter:	none

Response:	returns the alias

###### :PHYSical?

This command is used to query the physical address defined by the hardware. It applies only to the FLM-0121, FLM-0125 and FLM-0141 boards.

Syntax:	:SYSTem:COMMunicate:ADDRess:PHYSical?

Parameter:	none

Response:	returns the physical address

###### :STATe <b>

This command is used to specify the addressing state. The use of this setting varies depending on the controller type. On the FLM-0121 and FLM-0125 boards, addressing is permanently enabled. On the FLM-0141 board, addressing is normally enabled, but can be disabled to allow communications with a single module using the standard, non-addressed syntax. On all other boards, addressing is normally disabled, but can be enabled to allow aliased addressing to be used.

Syntax:	:SYSTem:COMMunicate:ADDRess:STATe <b>

Parameters:	<b>	addressing state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the addressing state.

Syntax:	:SYSTem:COMMunicate:ADDRess:STATe?

Parameters:	none

Response:	returns the addressing state	0 = disabled; 1 = enabled

#### :FORWard Commands

This group of commands is used to set up and query the message forwarding settings. Message forwarding is used to pass commands and responses between the host computer and a remote device (e.g. an optical amplifier) attached to Serial Port 1. Note that this facility is only available on the FLM-0107, FLM-0130, FLM-0151 and FLM-0177 boards. For further information on setting up the serial port and using message forwarding, refer to Sections 5.1.3.4 and 5.3, respectively.

###### :STATe <b>

This command is used to specify the message forwarding state.

Syntax:	:SYSTem:COMMunicate:FORWard:STATe <b>

Parameters:	<b>	forwarding state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the message forwarding state.

Syntax:	:SYSTem:COMMunicate:FORWard:STATe?

Parameters:	none

Response:	returns the forwarding state	0 = disabled; 1 = enabled

###### :TERMinator<0|1> <i>

This command is used to specify the terminators used when forwarding messages to the remote device. Up to two terminator characters can be defined in terms of their ASCII codes. The terminator characters are referenced individually using a suffix to the command header.

Syntax:	:SYSTem:COMMunicate:FORWard:TERMinator<0|1> <i>

Parameters:	<i>	ASCII terminator	e.g. 13 (CR), 10 (LF), 0 (none)

Configured:	yes

Response:	none

###### : TERMinator<0|1>?

This command is used to query the terminators used when forwarding messages to the remote device.

Syntax:	:SYSTem:COMMunicate:FORWard:TERMinator<0|1>?

Parameters:	none

Response:	returns the ASCII terminator

#### :SERial<0> Commands

This group of commands is used to set up and query the settings for Serial Port 0. The port number is specified as a suffix to the command header. Note that there should be no space between the header and the port number. The commands are only available on the FLM-0121, FLM-0125 and FLM-0141 boards.

###### :MODE <i>

This command is used to specify the mode.

Syntax:	:SYSTem:COMMunicate:SERial<0>:MODE <i>

Parameters:	<i>	mode	0 (normal), 1 (listen only)

Configured:	yes

Response:	none

###### :MODE?

This command is used to query the mode.

Syntax:	:SYSTem:COMMunicate:SERial<0>:MODE?

Parameters:	none

Response:	returns the mode

#### :SERial<1> Commands

This group of commands is used to set up and query the serial port settings. The port number is specified as a suffix to the command header. Note that there should be no space between the header and the port number.

These commands are only available on FLM-0107, FLM-0130, FLM-0151 and FLM-0177 boards and allow Serial Port 1 to be set up to communicate with a remote device. Note that any changes made will not come into effect until the system settings have been saved and the controller reset. For further information on communicating with a remote device, refer to Sections 5.1.3.2 and 5.3.

Note that the settings for Serial Port 0 – the port used for communications with the host computer – cannot be changed.

###### :BAUD <i>

This command is used to specify the baud rate.

Syntax:	:SYSTem:COMMunicate:SERial<1>:BAUD <i>

Parameters:	<i>	baud rate	9600, 19200, 38400 or 57600

Configured:	yes

Response:	none

###### :BAUD?

This command is used to query the baud rate.

Syntax:	:SYSTem:COMMunicate:SERial<1>:BAUD?

Parameters:	none

Response:	returns the baud rate

###### :BITS <i>

This command is used to specify the number of data bits.

Syntax:	:SYSTem:COMMunicate:SERial<1>:BITS <i>

Parameters:	<i>	number of data bits	1 to 8

Configured:	yes

Response:	none

###### :BITS?

This command is used to query the number of data bits.

Syntax:	:SYSTem:COMMunicate:SERial<1>:BITS?

Parameters:	none

Response:	returns the number of data bits

###### :PARity <i>

This command is used to specify the parity.

Syntax:	:SYSTem:COMMunicate:SERial<1>:PARity <i>

Parameters:	<i>	parity	0 (none), 1 (odd) or 2 (even)

Configured:	yes

Response:	none

###### :PARity?

This command is used to query the parity.

Syntax:	:SYSTem:COMMunicate:SERial<1>:PARity?

Parameters:	none

Response:	returns the parity

###### :SBITs <i>

This command is used to specify the number of stop bits.

Syntax:	:SYSTem:COMMunicate:SERial<1>:SBITs <i>

Parameters:	<i>	number of stop bits	1 or 2

Configured:	yes

Response:	none

###### :SBITs?

This command is used to query the number of stop bits.

Syntax:	:SYSTem:COMMunicate:SERial<1>:SBITs?

Parameters:	none

Response:	returns the number of stop bits

### :ERRor Commands

This group of commands is used to manage the error queue. As errors occur they are placed into a queue, which can store up to 10 error messages on a first-in, first-out (FIFO) basis. Querying the queue reads the oldest message and removes it from the queue.

Each error message comprises an error code and a brief description, separated by a comma. Error codes are in the range -32768 to +32767. Negative codes are defined by the SCPI standard; positive codes are used for errors specific to the controller. The error codes are listed in Table 5.38.

When there is no error, the message 0,“No error” is placed into the first location in the error queue. If the queue becomes full, the message -350,“Queue Overflow” is placed into the last location to indicate that an error (or errors) occurred but could not be recorded. The error queue may be cleared using the *CLS command (see Section 4.1).

###### :ERRor?

This command is used to query the error queue. It reads the next (oldest) error and removes it from the queue.

Syntax:	:SYSTem:ERRor?

Parameter:	none

Response:	returns the next error from the error queue

### :FIRMware Commands

###### :DATE?

This command is used to query the release date of the controller’s firmware.

Syntax:	:SYSTem:FIRMware:DATE?

Parameter:	none

Response:	returns the release date in the format <y>,<m>,<d>

###### :VERSion?

This command is used to query the version of the controller’s firmware.

Syntax:	:SYSTem:FIRMware:VERSion?

Parameter:	none

Response:	returns the version as a string

### :HARDware Commands

###### :APPLication <s>

This command is used to specify the controller’s application.

Syntax:	:SYSTem:HARDware:APPLication <s>

Parameter:	<s>	application as a string of up to 15 characters

Configured:	yes

Response:	none

###### :APPLication?

This command is used to query the controller’s application.

Syntax:	:SYSTem:HARDware:APPLication?

Parameter:	none

Response:	returns the application as a string

###### :DATE <y>,<m>,<d>

This command is used to specify the controller’s manufacture date.

Syntax:	:SYSTem:HARDware:DATE <y>,<m>,<d>

Parameters:	<y>	the year, expressed as a 4-digit integer

<m>	the month, expressed as an integer in the range 1 to 12

<d>	the day, expressed as an integer in the range 1 to 31

Configured:	yes

Response:	none

###### :DATE?

This command is used to query the controller’s manufacture date.

Syntax:	:SYSTem:HARDware:DATE?

Parameter:	none

Response:	returns the manufacture date in the format <y>,<m>,<d>

###### :LCDWelcome <s>

This command is used to specify the controller’s LCD welcome message. It allows a custom welcome message to be set into the FLM-0107 (Rev. B onwards), FLM-0130, FLM-0151 and FLM-0177 boards, when used in conjunction with an LCD module. The welcome message comprises 2 lines of 16 characters, but note that these should be concatenated into a single 32-character string when sending this command.

Syntax:	:SYSTem:HARDware:LCDWelcome <s>

Parameter:	<s>	LCD welcome message as a string of 32 characters

Configured:	yes

Response:	none

###### :LCDWelcome?

This command is used to query the controller’s LCD welcome message.

Syntax:	:SYSTem:HARDware:LCDWelcome?

Parameter:	none

Response:	returns the LCD welcome message as a string

###### :MANufacturer <s>

This command is used to specify the controller’s manufacturer.

Syntax:	:SYSTem:HARDware:MANufacturer <s>

Parameter:	<s>	manufacturer as a string of up to 29 characters

Configured:	yes

Response:	none

###### :MANufacturer?

This command is used to query the controller’s manufacturer.

Syntax:	:SYSTem:HARDware:MANufacturer?

Parameter:	none

Response:	returns the manufacturer as a string

###### :MODel <s>

This command is used to specify the controller’s model number.

Syntax:	:SYSTem:HARDware:MODel <s>

Parameter:	<s>	model number as a string of up to 14 characters

Configured:	yes

Response:	none

###### :MODel?

This command is used to query the controller’s model number.

Syntax:	:SYSTem:HARDware:MODel?

Parameter:	none

Response:	returns the model number as a string

###### :OPTions <i>

This command is used to specify the PCB options. It allows the number of pumps and the temperature controller types to be configured, along with various other options depending on board type and firmware version.

Syntax:	:SYSTem:HARDware:OPTions <i>

Parameter:	<i>	PCB options	see Table 5.11

Configured:	yes

Response:	none

###### :OPTions?

This command is used to query the PCB options.

Syntax:	:SYSTem:HARDware:OPTions?

Parameter:	none

Response:	returns the PCB options	see Table 5.11

Table 5.11 – PCB Options

| Bit | Description |
| --- | --- |
| 0 | Pump 0 (always 1 = fitted) |
| 1 | Pump 1 (0 = not fitted; 1 = fitted) |
| 2 | Temp. Control 0 type (always 0 = TEC) |
| 3 | Temp. Control 1 type (always 0 = TEC) |
| 4 | Temp. Control 2 type (0 = TEC; 1 = Heater) |
| 5 | Temp. Control 3 type (0 = TEC; 1 = Heater) |
| 6 | High Power option (0 = not fitted; 1 = fitted) |
| 7 | No Piezo option (0= piezo; 1 = no piezo) |
| 8 | Wideband Detector (0 = not fitted; 1 = fitted) |
| 9 | Variable Optical Attenuator (0 = not fitted; 1 = fitted) |
| 10 | High Gain Detector (0 = not fitted; 1 = fitted) |

Note:	High Power option requires firmware version 4.3.000 onwards.

No Piezo option requires firmware version 5.2.002 onwards and applies to FLM-0130, FLM-0151 & FLM-0177 PCB’s only.

Wideband Detector option requires firmware version 5.3.000 onwards and applies to FLM-0107, FLM-0130, FLM-0151 & FLM-0177 PCB’s only.

Variable Optical Attenuator and High Gain Detector options require firmware version 7.0.000 onwards and apply to FLM-0151 & FLM-0177 PCB’s only.

###### :REVision <a>

This command is used to specify the PCB revision.

Syntax:	:SYSTem:HARDware:REVision <a>

Parameter:	<a>	PCB revision

Configured:	yes

Response:	none

###### :REVision?

This command is used to query the PCB revision.

Syntax:	:SYSTem:HARDware:REVision?

Parameter:	none

Response:	returns the PCB revision

###### :SERial <s>

This command is used to specify the controller’s serial number.

Syntax:	:SYSTem:HARDware:SERial <s>

Parameter:	<s>	serial number as a string of up to 8 characters

Configured:	yes

Response:	none

###### :SERial?

This command is used to query the controller’s serial number.

Syntax:	:SYSTem:HARDware:SERial?

Parameter:	none

Response:	returns the serial number as a string

###### :TYPE <i>

This command is used to specify the PCB type. It also resets the configured settings to the defaults for the PCB type specified.  For this reason, the command uses a special password protection mechanism and is reserved for factory use only.

Syntax:	:SYSTem:HARDware:TYPE <i>

Parameter:	<i>	PCB type	see Table 5.12

Configured:	yes

Response:	none

###### :TYPE?

This command is used to query the PCB type.

Syntax:	:SYSTem:HARDware:TYPE?

Parameter:	none

Response:	returns the PCB type	see Table 5.12

Table 5.12 – PCB Types

| Type | Description |
| --- | --- |
| 100 | FLM-0100 Fiber Laser Controller (single-pump) |
| 103 | FLM-0103 Fiber Laser Controller (twin-pump) |
| 105 | FLM-0105 ASE Controller |
| 106 | FLM-0106 Fiber Laser Controller (single-pump) |
| 107 | FLM-0107 Fiber Laser Controller (twin-pump) |
| 121 | FLM-0121 Fiber Laser Controller (single-pump) |
| 125 | FLM-0125 Fiber Laser Controller (single-pump) |
| 130 | FLM-0130 Fiber Laser Controller (twin-pump) |
| 141 | FLM-0141 Fiber Laser Controller (twin-pump) |
| 151 | FLM-0151 Fiber Laser Controller (twin-pump) |
| 177 | FLM-0177 Fiber Laser Controller (twin-pump) |
| 0 | Restores default settings on current PCB type. |

Note:	FLM-0105 requires firmware version 3.0.000 onwards.

FLM-0106 requires firmware version 4.0.000 onwards.

FLM-0107 requires firmware version 4.0.000 onwards.

FLM-0121 requires firmware version 5.0.000 onwards.

FLM-0125 requires firmware version 5.0.000 onwards.

FLM-0130 requires firmware version 5.0.000 onwards.

FLM-0141 requires firmware version 6.0.000 onwards.

FLM-0151 requires firmware version 7.0.000 onwards.

FLM-0177 requires firmware version 8.0.000 onwards.

Type 0 requires firmware version 6.5.000 onwards.

### :MHFTuning Commands

This group of commands is used to set up and query the mode hop free (MHF) tuning settings. Note that they are not available on the FLM-0105 ASE Controller.

###### :NBSLope <f>

This command is used to specify the temperature sensitivity of the narrow-band grating.

Syntax:	:SYSTem:MHFTuning:NBSLope <f>	(f/w version 1.0.001 to 4.9.005)

Parameters:	<f>	narrow-band temperature sensitivity in pm/ºC

Configured:	yes

Response:	none

###### :NBSLope?

This command is used to query the temperature sensitivity of the narrow-band grating.

Syntax:	:SYSTem:MHFTuning:NBSLope?	(f/w version 1.0.001 to 4.9.005)

Parameters:	none

Response:	returns the narrow-band temperature sensitivity in pm/ºC

###### :WBSLope <f>

This command is used to specify the temperature sensitivity of the wide-band grating.

Syntax:	:SYSTem:MHFTuning:WBSLope <f>	(f/w version 1.0.001 to 4.9.005)

Parameters:	<f>	wide-band temperature sensitivity in pm/ºC

Configured:	yes

Response:	none

###### :WBSLope?

This command is used to query the temperature sensitivity of the wide-band grating.

Syntax:	:SYSTem:MHFTuning:WBSLope?	(f/w version 1.0.001 to 4.9.005)

Parameters:	none

Response:	returns the wide-band temperature sensitivity in pm/ºC

#### :BAND<a|i> Commands

This group of commands is used to set up and query the settings for each MHF band. The band is specified using either a letter or an integer suffix to the command header, depending on the firmware version. The valid ranges are A - Z (or a - z) and 0 - 59 respectively. Note that there should be no space between the header and the band suffix.

The wavelength calibration table comprises one or more defined bands. Bands are filled sequentially, starting from the first (A or 0). Each band is defined in terms of its centre, or reference, wavelength and its half-width. Additional parameters are used to define the conditons needed to control the wavelength within the band. Normally, only a few bands are defined and the end of data is marked using a zero value in the reference wavelength field of the next band, i.e. the first undefined band has its reference wavelength set to zero.

For firmware version 5.0.000 onwards, the total wavelength range can be determined as follows:

minimum wavelength   =  reference wavelengthBand 0  –  half-widthBand 0

maximum wavelength  =  reference wavelengthBand n  +  half-widthBand n

where Band 0 and Band n are, respectively, the first and last defined bands.

###### :HALFwidth <f>

This command is used to specify the half-width of the band.

Syntax:	:SYSTem:MHFTuning:BAND<i>:HALFwidth <f>	(f/w version 5.0.000 onwards)

Parameters:	<f>	half-width in pm

Configured:	yes

Response:	none

###### :HALFwidth?

This command is used to query the half-width of the band.

Syntax:	:SYSTem:MHFTuning:BAND<i>:HALFwidth?	(f/w version 5.0.000 onwards)

Parameters:	none

Response:	returns the half-width in pm

###### :NBSLope <f>

This command is used to specify the temperature sensitivity of the narrow-band grating on the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards.

Syntax:	:SYSTem:MHFTuning:BAND<i>:NBSLope <f>	(f/w version 5.0.000 onwards)

Parameters:	<f>	narrow-band temperature sensitivity in pm/ºC

Configured:	yes

Response:	none

###### :NBSLope?

This command is used to query the temperature sensitivity of the narrow-band grating on the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards.

Syntax:	:SYSTem:MHFTuning:BAND<i>:NBSLope?	(f/w version 5.0.000 onwards)

Parameters:	none

Response:	returns the narrow-band temperature sensitivity in pm/ºC

###### :NBTemp <f>

This command is used to specify the temperature of the narrow-band grating at the center of the band. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<a>:NBTemp <f>	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:NBTemp <f>	(f/w version 4.1.000 onwards)

Parameters:	<f>	narrow-band temperature in ºC

Configured:	yes

Response:	none

###### :NBTemp?

This command is used to query the temperature of the narrow-band grating at the center of the band. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<a>:NBTemp?	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:NBTemp?	(f/w version 4.1.000 onwards)

Parameters:	none

Response:	returns the narrow-band temperature in ºC

###### :RTEMperature <f>

This command is used to specify the heater temperature at the center of the band. It applies to the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<i>:RTEMperature <f>	(f/w version 5.0.000 onwards)

Parameters:	<f>	reference temperature in °C

Configured:	yes

Response:	none

###### :RTEMperature?

This command is used to query the heater temperature at the center of the band. It applies to the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<i>:RTEMperature?	(f/w version 5.0.000 onwards)

Parameters:	none

Response:	returns the reference temperature in °C

###### :RVOLtage <f>

This command is used to specify the piezo drive voltage at the center of the band. It applies to the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<i>:RVOLtage <f>	(f/w version 5.0.000 onwards)

Parameters:	<f>	reference voltage in V

Configured:	yes

Response:	none

###### :RVOLtage?

This command is used to query the piezo drive voltage at the center of the band. It applies to the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<i>:RVOLtage?	(f/w version 5.0.000 onwards)

Parameters:	none

Response:	returns the reference voltage in V

###### :RWAVelegth <f>

This command is used to specify the wavelength at the center of the band.

Syntax:	:SYSTem:MHFTuning:BAND<a>:RWAVelength <f>	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:RWAVelength <f>	(f/w version 4.1.000 onwards)

Parameters:	<f>	reference wavelength in nm

Configured:	yes

Response:	none

###### :RWAVelegth?

This command is used to query the wavelength at the center of the band.

Syntax:	:SYSTem:MHFTuning:BAND<a>:RWAVelength?	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:RWAVelength?	(f/w version 4.1.000 onwards)

Parameters:	none

Response:	returns the reference wavelength in nm

###### :TSLope <f>

This command is used to specify the temperature sensitivity on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards.

Syntax:	:SYSTem:MHFTuning:BAND<i>:TSLope <f>	(f/w version 5.0.000 onwards)

Parameters:	<f>	temperature sensitivity in pm/ºC

Configured:	yes

Response:	none

###### :TSLope?

This command is used to query the temperature sensitivity on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards.

Syntax:	:SYSTem:MHFTuning:BAND<i>:TSLope?	(f/w version 5.0.000 onwards)

Parameters:	none

Response:	returns the temperature sensitivity in pm/ºC

###### :VSLope <f>

This command is used to specify the voltage sensitivity on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards.

Syntax:	:SYSTem:MHFTuning:BAND<i>:VSLope <f>	(f/w version 5.0.000 onwards)

Parameters:	<f>	voltage sensitivity in pm/V

Configured:	yes

Response:	none

###### :VSLope?

This command is used to query the voltage sensitivity on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-151 and FLM-177 boards.

Syntax:	:SYSTem:MHFTuning:BAND<i>:VSLope?	(f/w version 5.0.000 onwards)

Parameters:	none

Response:	returns the voltage sensitivity in pm/V

###### :WBHalfrange <f>

This command is used to specify the half-range of the wide-band grating for the given band. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<a>:WBHalfrange <f>	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:WBHalfrange <f>	(f/w version 4.1.000 to 4.9.005)

Parameters:	<f>	wide-band half-range in ºC

Configured:	yes

Response:	none

###### :WBHalfrange?

This command is used to query the half-range of the wide-band grating for the given band. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<a>:WBHalfrange?	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:WBHalfrange?	(f/w version 4.1.000 to 4.9.005)

Parameters:	none

Response:	returns the wide-band half-range in ºC

###### :WBSLope <f>

This command is used to specify the temperature sensitivity of the wide-band grating. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<i>:WBSLope <f>	(f/w version 5.0.000 onwards)

Parameters:	<f>	wide-band temperature sensitivity in pm/ºC

Configured:	yes

Response:	none

###### :WBSLope?

This command is used to query the temperature sensitivity of the wide-band grating on the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards.

Syntax:	:SYSTem:MHFTuning:BAND<i>:WBSLope?	(f/w version 5.0.000 onwards)

Parameters:	none

Response:	returns the wide-band temperature sensitivity in pm/ºC

###### :WBTemp <f>

This command is used to specify the temperature of the wide-band grating at the center of the band. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<a>:WBTemp <f>	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:WBTemp <f>	(f/w version 4.1.000 onwards)

Parameters:	<f>	wide-band temperature in ºC

Configured:	yes

Response:	none

###### :WBTemp?

This command is used to query the temperature of the wide-band grating at the center of the band. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:MHFTuning:BAND<a>:WBTemp?	(f/w version 1.0.001 to 4.0.002)

:SYSTem:MHFTuning:BAND<i>:WBTemp?	(f/w version 4.1.000 onwards)

Parameters:	none

Response:	returns the wide-band temperature in ºC

### :OPTimization Commands

This group of commands is used to set up and query the optimization settings. Note that they are only available on the FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards.

###### :CONFidence <i>

This command is used to specify the minimum confidence level. The optimization voltage will only be applied if the confidence flag is greater than or equal to this value. For further details, refer to the :OPTimization:FLAG? command in Section 5.5.

Syntax:	:SYSTem:OPTimization:CONFidence <i>

Parameters:	<i>	confidence level	0 (always apply), 1, 2, 3 or 4 (never apply)

Configured:	yes

Response:	none

###### :CONFidence?

This command is used to query the minimum confidence level.

Syntax:	:SYSTem:OPTimization:CONFidence?

Parameters:	none

Response:	returns the minimum confidence level

###### :SAVE <b>

This command is used to specify the auto-save state. When auto-save is enabled, the optimization voltage will be stored in non-volatile memory.

Syntax:	:SYSTem:OPTimization:SAVE <b>

Parameters:	<b>	auto-save state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :SAVE?

This command is used to query the auto-save state.

Syntax:	:SYSTem:OPTimization:SAVE?

Parameters:	none

Response:	returns the auto-save state

#### :DIP Commands

This group of commands is used to set up and query the dip detection settings used when analyzing the ramp data.

###### :DMIN <f>

This command is used to specify the minimum depth of a power dip. The value is specified as a proportion of the total power range observed during the ramp.

Syntax:	:SYSTem:OPTimization:DIP:DMIN <f>

Parameters:	<f>	minimum depth

Configured:	yes

Response:	none

###### :DMIN?

This command is used to query the minimum depth of a power dip.

Syntax:	:SYSTem:OPTimization:DIP:DMIN?

Parameters:	none

Response:	returns the minimum depth

###### :SMIN <f>

This command is used to specify the minimum spacing between power dips.

Syntax:	:SYSTem:OPTimization:DIP:SMIN <f>

Parameters:	<f>	minimum spacing in V

Configured:	yes

Response:	none

###### :SMIN?

This command is used to query the minimum spacing between power dips.

Syntax:	:SYSTem:OPTimization:DIP:SMIN?

Parameters:	none

Response:	returns the minimum spacing in V

###### :STATe <b>

This command is used to specify the dip detection state.

Syntax:	:SYSTem:OPTimization:DIP:STATe <b>

Parameters:	<b>	dip detection state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the dip detection state.

Syntax:	:SYSTem:OPTimization:DIP:STATe?

Parameters:	none

Response:	returns the dip detection state	0 = disabled; 1 = enabled

###### :WMAX <f>

This command is used to specify the maximum width of a power dip.

Syntax:	:SYSTem:OPTimization:DIP:WMAX <f>

Parameters:	<f>	maximum width in V

Configured:	yes

Response:	none

###### :WMAX?

This command is used to query the maximum width of a power dip.

Syntax:	:SYSTem:OPTimization:DIP:WMAX?

Parameters:	none

Response:	returns the maximum width in V

#### :PEAK Commands

This group of commands is used to set up and query the peak location settings used when analyzing the ramp data.

###### :PMAX<0|1> <f>

This command is used to specify the maximum peak voltages. Note that separate values are used for the upward (PMAX0) and downward (PMAX1) ramps.

Syntax:	:SYSTem:OPTimization:PEAK:PMAX<0|1> <f>

Parameters:	<f>	maximum voltage in V

Configured:	yes

Response:	none

###### :PMAX<0|1>?

This command is used to query the maximum peak voltages.

Syntax:	:SYSTem:OPTimization:PEAK:PMAX<0|1>?

Parameters:	none

Response:	returns the maximum voltage in V

###### :PMIN<0|1> <f>

This command is used to specify the minimum peak voltages. Note that separate values are used for the upward (PMIN0) and downward (PMIN1) ramps.

Syntax:	:SYSTem:OPTimization:PEAK:PMIN<0|1> <f>

Parameters:	<f>	minimum voltage in V

Configured:	yes

Response:	none

###### :PMIN<0|1>?

This command is used to query the minimum peak voltages.

Syntax:	:SYSTem:OPTimization:PEAK:PMIN<0|1>?

Parameters:	none

Response:	returns the minimum voltage in V

###### :SMAX <f>

This command is used to specify the maximum peak separation.

Syntax:	:SYSTem:OPTimization:PEAK:SMAX <f>

Parameters:	<f>	maximum separation in V

Configured:	yes

Response:	none

###### :SMAX?

This command is used to query the maximum peak separation.

Syntax:	:SYSTem:OPTimization:PEAK:SMAX?

Parameters:	none

Response:	returns the maximum separation in V

###### :SMIN <f>

This command is used to specify the minimum peak separation.

Syntax:	:SYSTem:OPTimization:PEAK:SMIN <f>

Parameters:	<f>	minimum separation in V

Configured:	yes

Response:	none

###### :SMIN?

This command is used to query the minimum peak separation.

Syntax:	:SYSTem:OPTimization:PEAK:SMIN?

Parameters:	none

Response:	returns the minimum separation in V

#### :QUADratic Commands

This group of commands is used to set up and query the quadratic fit settings used when analyzing the ramp data.

###### :CMAX <f>

This command is used to specify the maximum limit on the 2nd order quadratic coefficient.

Syntax:	:SYSTem:OPTimization:QUADratic:CMAX <f>

Parameters:	<f>	maximum 2nd order coefficient

Configured:	yes

Response:	none

###### :CMAX?

This command is used to query the maximum limit on the 2nd order quadratic coefficient.

Syntax:	:SYSTem:OPTimization:QUADratic:CMAX?

Parameters:	none

Response:	returns the maximum 2nd order coefficient

###### :CMIN <f>

This command is used to specify the minimum limit on the 2nd order quadratic coefficient.

Syntax:	:SYSTem:OPTimization:QUADratic:CMIN <f>

Parameters:	<f>	minimum 2nd order coefficient

Configured:	yes

Response:	none

###### :CMIN?

This command is used to query the minimum limit on the 2nd order quadratic coefficient.

Syntax:	:SYSTem:OPTimization:QUADratic:CMIN?

Parameters:	none

Response:	returns the minimum 2nd order coefficient

###### :NMIN <i>

This command is used to specify the minimum number of points required for a valid quadratic fit.

Syntax:	:SYSTem:OPTimization:QUADratic:NMIN <i>

Parameters:	<i>	minimum number of points

Configured:	yes

Response:	none

###### :NMIN?

This command is used to query the minimum number of points required for a valid quadratic fit.

Syntax:	:SYSTem:OPTimization:QUADratic:NMIN?

Parameters:	none

Response:	returns the minimum number of points

###### :RMIN <f>

This command is used to specify the minimum limit on the coefficient of determination. The coefficient of determination provides a measure of the quadratic fit quality. For further details, refer to the :OPTimization:RAMP<0|1>QUADratic<0|1>? command in Section 5.5.2.

Syntax:	:SYSTem:OPTimization:QUADratic:RMIN <f>

Parameters:	<f>	minimum coefficient of determination

Configured:	yes

Response:	none

###### :RMIN?

This command is used to query the minimum limit on the coefficient of determination.

Syntax:	:SYSTem:OPTimization:QUADratic:RMIN?

Parameters:	none

Response:	returns the minimum coefficient of determination

#### :RAMP Commands

This group of commands is used to set up and query the ramp settings.

###### :IDWell <i>

This command is used to specify the initial dwell time at the start of the ramp.

Syntax:	:SYSTem:OPTimization:RAMP:IDWell <i>

Parameters:	<i>	initial dwell time in s/10

Configured:	yes

Response:	none

###### :IDWell?

This command is used to query the initial dwell time at the start of the ramp.

Syntax:	:SYSTem:OPTimization:RAMP:IDWell?

Parameters:	none

Response:	returns the initial dwell time in s/10

###### :SDWell <i>

This command is used to specify the dwell time at each step on the ramp.

Syntax:	:SYSTem:OPTimization:RAMP:SDWell <i>

Parameters:	<i>	step dwell time in s/10

Configured:	yes

Response:	none

###### :SDWell?

This command is used to query the dwell time at each step on the ramp.

Syntax:	:SYSTem:OPTimization:RAMP:SDWell?

Parameters:	none

Response:	returns the step dwell time in s/10

###### :STEPs <i>

This command is used to specify the number of steps on the ramp.

Syntax:	:SYSTem:OPTimization:RAMP:STEPs <i>

Parameters:	<i>	number of steps

Configured:	yes

Response:	none

###### :STEPs?

This command is used to query the number of steps on the ramp.

Syntax:	:SYSTem:OPTimization:RAMP:STEPs?

Parameters:	none

Response:	returns the number of steps

###### :VMAX <f>

This command is used to specify the maximum ramp voltage.

Syntax:	:SYSTem:OPTimization:RAMP:VMAX <f>

Parameters:	<f>	maximum voltage in V

Configured:	yes

Response:	none

###### :VMAX?

This command is used to query the maximum ramp voltage.

Syntax:	:SYSTem:OPTimization:RAMP:VMAX?

Parameters:	none

Response:	returns the maximum voltage in V

###### :VMIN <f>

This command is used to specify the minimum ramp voltage.

Syntax:	:SYSTem:OPTimization:RAMP:VMIN <f>

Parameters:	<f>	minimum voltage in V

Configured:	yes

Response:	none

###### :VMIN?

This command is used to query the minimum ramp voltage.

Syntax:	:SYSTem:OPTimization:RAMP:VMIN?

Parameters:	none

Response:	returns the minimum voltage in V

### :PADJust Commands

This group of commands is used to set up and query the power adjust settings. They allow the laser power to be adjusted via a potentiometer input, a BNC input or from the command interface. Note that these commands are only available on the FLM-0107 (Rev. B onwards), FLM-0130, FLM-0151 and FLM-0177 boards, when configured as a high power source, and also on the FLM-0151 and FLM-0177 boards, when fitted with the variable optical attenuator option.

###### :INPut <i>

This command is used to specify the analog input to be used for adjusting the power. When the potentiometer input is selected, the firmware compensates to give the control knob a linear response. Note that this facitlity only applies to the Gen. II and Gen. III high power sources (hardware application = “HPS-K” or “HPS-I”).

Syntax:	:SYSTem:PADJust:INPut <i>

Parameters:	<i>	input	0 = BNC socket; 1 = potentiometer

Configured:	yes

Response:	none

###### :INPut?

This command is used to query the analog input used for adjusting the power.

Syntax:	:SYSTem:PADJust:INPut?

Parameters:	none

Response:	returns the input

###### :PMAX <f>

This command is used to specify the maximum power. It is used in conjunction with the :SYSTem:PADJust:PMIN <f> command to set up the range over which the power can be adjusted. This feature is only available on the Gen. II high power source (hardware application = “HPS-K”).

Syntax:	:SYSTem:PADJust:PMAX <f>

Parameters:	<f>	maximum power in mW

Configured:	yes

Response:	none

###### :PMAX?

This command is used to query the maximum power.

Syntax:	:SYSTem:PADJust:PMAX?

Parameters:	none

Response:	returns the maximum power in mW

###### :PMIN <f>

This command is used to specify the minimum power. It is used in conjunction with the :SYSTem:PADJust:PMAX <f> command to set up the range over which the power can be adjusted. This feature is only available on the Gen. II high power source (hardware application = “HPS-K”).

Syntax:	:SYSTem:PADJust:PMIN <f>

Parameters:	<f>	minimum power in mW

Configured:	yes

Response:	none

###### :PMIN?

This command is used to query the minimum power.

Syntax:	:SYSTem:PADJust:PMIN?

Parameters:	none

Response:	returns the minimum power in mW

###### :POWer <f>

This command is used to specify the power when the analog inputs are not in use (i.e. when the power adjust state is set to 0). This feature is only available on the Gen. II high power source (hardware application = “HPS-K”) and on the FLM-0151 and FLM-0177 boards, when fitted with the variable optical attenuator option.

Syntax:	:SYSTem:PADJust:POWer <f>

Parameters:	<f>	power in mW

Configured:	yes

Response:	none

###### :POWer?

This command is used to query the power. It can be used at any time, irrespective of whether the power is being set via the potentiometer input, the BNC input or the command interface.

Syntax:	:SYSTem:PADJust:POWer?

Parameters:	none

Response:	returns the power in mW

###### :POWer:MUTE <b>

This command is used to specify the power mute state. It is only available on FLM-0151 and FLM-0177 boards, when fitted with the variable optical attenuator option.

Syntax:	:SYSTem:PADJust:POWer:MUTE <b>

Parameters:	<b>	mute state	0 = disabled; 1 = enabled

Configured:	no

Response:	none

###### :POWer:MUTE?

This command is used to query the power mute state. It is only available on FLM-0151 and FLM-0177 boards, when fitted with the variable optical attenuator option.

Syntax:	:SYSTem:PADJust:POWer:MUTE?

Parameters:	none

Response:	returns the mute state	0 = disabled; 1 = enabled

###### :STATe <i>

This command is used to specify the adjustment state. When the adjustment state is set to 1, the potentiometer (or BNC) input is sampled and used to set the power. When the adjustment state is set to 0, the potentiometer (or BNC) input is disabled. Note that this facility only applies to the Gen. II, Gen. III and Gen. IV high power sources (hardware application = “HPS-K”, “HPS-I” and “HPS-4”).

Syntax:	:SYSTem:PADJust:STATe <i>

Parameters:	<i>	adjustment state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the adjustment state.

Syntax:	:SYSTem:PADJust:STATe?

Parameters:	none

Response:	returns the adjustment state

### :PASSword Commands

This group of commands is used to enable and disable the commands associated with the controller’s configured data set. During normal operation, the controller’s configuration (serial number, manufacture date, pump control settings, temperature control settings, etc.) is protected to guard against any vital parameters being modified inadvertently. However, during manufacture, when these parameters are being set up, the associated commands must be enabled.

The configurable parameters protected by the password mechanism are identified in the Config. column of Table 5.33. The *SAV command (see Section 4.5), which is used to save the controller’s configuration into non-volatile memory, is also protected.

Note that the controller always starts up in the protected state, both at switch-on and following a software reset (*RST) command.

Passwords are not case-sensitive. The default password is “NP”.

###### :CDISable <s>

This command is used to disable the protected commands.

Syntax:	:SYSTem:PASSword:CDISable <s>

Parameter:	<s>	password as a string of up to 15 characters

Configured:	no

Response:	none

###### :CENable <s>

This command is used to enable the protected commands.

Syntax:	:SYSTem:PASSword:CENable <s>

Parameter:	<s>	password as a string of up to 15 characters

Configured:	no

Response:	none

###### :CENable:STATe?

This command is used to query the state of the protected commands.

Syntax:	:SYSTem:PASSword:CENable:STATe?

Parameter:	none

Response:	returns the protection state	0 = disabled; 1 = enabled

###### :NEW <s1>,<s2>

This command is used to specify a new password.

Syntax:	:SYSTem:PASSword:NEW <s1>,<s2>

Parameter:	<s1>	old password as a string of up to 15 characters

<s2>	new password as a string of up to 15 characters

Response:	none

### :PCONtrol Commands

This group of commands is used to set up and query the pump control settings.

###### :CLIMit[0|1] <i>

This command is used to specify the current limit. On the twin-pump boards, a separate current limit setting is provided for each pump.

Syntax:	:SYSTem:PCONtrol:CLIMit <i>	FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PCONtrol:CLIMit <i>	FLM-0100/3/6	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1> <i>	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:CLIMit <i>	FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1> <i>	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1> <i>	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1> <i>	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1> <i>	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	<i>	current limit in mA

Configured:	yes

Response:	none

###### :CLIMit[0|1]?

This command is used to query the current limit.

Syntax:	:SYSTem:PCONtrol:CLIMit?	FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PCONtrol:CLIMit?	FLM-0100/3/6	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1>?	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:CLIMit?	FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1>?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1>?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1>?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PCONtrol:CLIMit<0|1>?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the current limit in mA

###### :DELay <b>

This command is used to specify the startup delay state.

Syntax:	:SYSTem:PCONtrol:DELay <b>

Parameters:	<b>	delay state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :DELay?

This command is used to query the startup delay state.

Syntax:	:SYSTem:PCONtrol:DELay?

Parameters:	none

Response:	returns the delay state	0 = disabled; 1 = enabled

###### :EMODulation <b>

This command is used to specify the external modulation state. It is available on the FLM-0151 board only.

Syntax:	:SYSTem:PCONtrol:EMODulation <b>

Parameters:	<b>	modulation state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :EMODulation?

This command is used to query the external modulation state. It is available on the FLM-0151 board only.

Syntax:	:SYSTem:PCONtrol:EMODulation?

Parameters:	none

Response:	returns the modulation state	0 = disabled; 1 = enabled

###### :FLAGs?

This command is used to query the pump control status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:PCONtrol:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.13

Table 5.13 – Pump Control Flags

| Bit | Description |
| --- | --- |
| 0 | Mode (0 = Power; 1 = Current) |
| 1 | Disabled (via jumper) |
| 2 | On |
| 3 | Over-current |
| 4 | Shutdown |
| 5 | Startup |

Note:	All bits are active high.

###### :MSELect <i>

This command is used to specify the operating mode.

Syntax:	:SYSTem:PCONtrol:MSELect <i>

Parameters:	<i>	mode select code	see Table 5.14

Configured:	yes

Response:	none

Table 5.14 – Pump Operating Modes

| Code | Description |
| --- | --- |
| 0 | Constant power |
| 1 | Constant current |
| 2 | Hardware select (via jumper) |
| 3 | Constant current + optimization |

Note:	Mode 2 not available on FLM-0105, FLM-0121, FLM-0125 & FLM-0141.

Mode 3 available on FLM-0130, FLM-0151 & FLM-0177 only and requires firmware version 5.4.008 (FLM-0130), 7.0.000 (FLM-0151) or 8.0.000 (FLM-0177) onwards.

###### :MSELect?

This command is used to query the operating mode.

Syntax:	:SYSTem:PCONtrol:MSELect?

Parameters:	none

Response:	returns the mode select code 	see Table 5.14

###### :PLIMit <i>

###### :PLIMit <f>

This command is used to specify the power limit.

Syntax:	:SYSTem:PCONtrol:PLIMit <i>	(f/w version 5.4.012 to 5.4.014)

:SYSTem:PCONtrol:PLIMit <f>	(f/w version 6.2.000 onwards)

Parameters:	<i>	power limit in mW		(f/w version 5.4.012 to 5.4.014)

<f>	power limit in mW		(f/w version 6.2.000 onwards)

Configured:	yes

Response:	none

###### :PLIMit?

This command is used to query the power limit.

Syntax:	:SYSTem:PCONtrol:PLIMit?	(f/w version 5.4.012 to 5.4.014)

:SYSTem:PCONtrol:PLIMit?	(f/w version 6.2.000 onwards)

Parameters:	none

Response:	returns the power limit in mW

###### :PLIMit:STATe <b>

This command is used to specify the power limiting state.

Syntax:	:SYSTem:PCONtrol:PLIMit:STATe <b>

Parameters:	<b>	power limiting state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :PLIMit:STATe?

This command is used to query the power limiting state.

Syntax:	:SYSTem:PCONtrol:PLIMit:STATe?

Parameters:	none

Response:	returns the power limiting state	0 = disabled; 1 = enabled

###### :PMODe <b>

This command is used to specify the pump power mode state. Note that this facility is only available on the FLM-0151 and FLM-0177 boards.

Syntax:	:SYSTem:PCONtrol:PMODe <b>

Parameters:	<b>	power mode state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :PMODe?

This command is used to query the pump power mode state. Note that this facility is only available on the FLM-0151 and FLM-0177 boards.

Syntax:	:SYSTem:PCONtrol:PMODe?

Parameters:	none

Response:	returns the power mode state	0 = disabled; 1 = enabled

###### :RRATe<0|1> <i>

This command is used to specify the pump set-point ramp rates on the Amplified ASE (hardware application = HPASE).

Syntax:	:SYSTem:PCONtrol:RRATe<0|1> <i>

Parameters:	<i>	current set-point ramp rate in mA/s/10

Configured:	yes

Response:	none

###### :RRATe<0|1>?

This command is used to query the pump set-point ramp rates on the Amplified ASE (hardware application = HPASE).

Syntax:	:SYSTem:PCONtrol:RRATe<0|1>?

Parameters:	none

Response:	returns	the current set-point ramp rate in mA/s/10

###### :SPOint <i>

###### :SPOint <f>

This command is used to specify the pump set-point. Note that it is only available in firmware versions 1.0.001 to 3.0.000.

In early versions of the firmware (see below), only the pump current can be set using this command; the later versions, however, allow either the power or current to be set, depending on the mode of operation at the time. When the controller is set for constant current operation (Mode 1), this command will also select the appropriate current range automatically.

The :SYSTem:PCONtrol:SPOint commands provide a higher-level, more user-friendly means of accessing the pump set-point than the :SYSTem:PCONtrol:MODE<0|1>:SPOint commands (see Section 5.1.11.1). In addition, they do not require the controller to be unlocked to adjust the set-point (though to save the new settings to non-volatile memory does). When full control of both the operating range and set-point is needed, use the :SYSTem:PCONtrol:MODE<0|1>:SPOint commands.

Syntax:	:SYSTem:PCONtrol:SPOint <i>	(f/w version 1.0.xxx series only)

:SYSTem:PCONtrol:SPOint <f>	(f/w version 1.1.000 to 3.0.000)

Parameters:	<i>	set-point in mA (Mode 1 only) 	(f/w version 1.0.xxx series only)

<f>	set-point in mW (Mode 0) or mA (Mode 1) 	(f/w version 1.1.000 to 3.0.000)

Configured:	yes

Response:	none

###### :SPOint?

This command is used to query the pump set-point.

Syntax:	:SYSTem:PCONtrol:SPOint?

Parameters:	none

Response:	returns the set-point in mA (Mode 1 only) 	(f/w version 1.0.xxx series only)

returns the set-point in mW (Mode 0) or mA (Mode 1) 	(f/w version 1.1.000 to 3.1.000)

###### :STATe <b>

This command is used to specify the pump state.

Syntax:	:SYSTem:PCONtrol:STATe <b>

Parameters:	<b>	pump state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the pump state.

Syntax:	:SYSTem:PCONtrol:STATe?

Parameters:	none

Response:	returns the pump state	0 = disabled; 1 = enabled

#### :MODE<0|1> Commands

This group of commands is used to set up and query the settings for each operating mode. The mode is specified as a suffix to the command header. Note that there should be no space between the header and the mode number.

###### :KD <f>

This command is used to specify the differential constant used by the power control loop.

Syntax:	:SYSTem:PCONtrol:MODE<0>:KD <f>

Parameters:	<f>	differential constant in bits/mW.s

Configured:	yes

Response:	none

###### :KD?

This command is used to query the differential constant used by the power control loop.

Syntax:	:SYSTem:PCONtrol:MODE<0>:KD?

Parameters:	none

Response:	returns the differential constant in bits/mW.s

###### :KI <f>

This command is used to specify the integral constant used by the power control loop.

Syntax:	:SYSTem:PCONtrol:MODE<0>:KI <f>

Parameters:	<f>	integral constant in bits/mW/s

Configured:	yes

Response:	none

###### :KI?

This command is used to query the integral constant used by the power control loop.

Syntax:	:SYSTem:PCONtrol:MODE<0>:KI?

Parameters:	none

Response:	returns the integral constant in bits/mW/s

###### :KP <f>

This command is used to specify the proportional constant used by the power control loop.

Syntax:	:SYSTem:PCONtrol:MODE<0>:KP <f>

Parameters:	<f>	proportional constant in bits/mW

Configured:	yes

Response:	none

###### :KP?

This command is used to query the proportional constant used by the power control loop.

Syntax:	:SYSTem:PCONtrol:MODE<0>:KP?

Parameters:	none

Response:	returns the proportional constant in bits/mW

###### :RANGe <i>

This command is used to specify the operating range used for the given mode.

Syntax:	:SYSTem:PCONtrol:MODE<0|1>:RANGe <i>

Parameters:	<i>	operating range	0, 1, 2 or 3	FLM-0100 & FLM-0103

0 only	FLM-0105, FLM-0106, FLM-0107,

FLM-0121, FLM-0125, FLM-0130,

FLM-0141, FLM-0151 & FLM-0177

Configured:	yes

Response:	none

###### :RANGe?

This command is used to query the operating range used for the given mode.

Syntax:	:SYSTem:PCONtrol:MODE<0|1>:RANGe?

Parameters:	none

Response:	returns the operating range	0, 1, 2 or 3	FLM-0100 & FLM-0103

0 only	FLM-0105, FLM-0106, FLM-0107,

FLM-0121, FLM-0125, FLM-0130,

FLM-0141, FLM-0151 & FLM-0177

###### :SPOint <i>

###### :SPOint[0|1] <f>

This command is used to specify the set-point for the given mode of operation. On the twin-pump boards, a separate set-point is provided for each pump.

Syntax:	:SYSTem:PCONtrol:MODE<0|1>:SPOint <i>	(f/w version 1.0.xxx series only)

:SYSTem:PCONtrol:MODE<0|1>:SPOint <f>	(f/w version 1.1.000 onwards)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1> <f>	(f/w ver. 4.0.000 onwards on FLM-0107)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1> <f>	(f/w ver. 5.0.000 onwards on FLM-0130)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1> <f>	(f/w ver. 6.0.000 onwards on FLM-0141)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1> <f>	(f/w ver. 7.0.000 onwards on FLM-0151)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1> <f>	(f/w ver. 8.0.000 onwards on FLM-0177)

Parameters:	<i>	set-point in bits (Mode 0 & Mode 1)	(f/w version 1.0.xxx series only)

<f>	set-point in mW (Mode 0) or mA (Mode 1)	(f/w version 1.1.000 onwards)

Configured:	yes

Response:	none

###### :SPOint[0|1]?

This command is used to query the set-point for the given mode of operation.

Syntax:	:SYSTem:PCONtrol:MODE<0|1>:SPOint?	(f/w version 1.0.001 onwards)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1>?	(f/w ver. 4.0.000 onwards on FLM-0107)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1>?	(f/w ver. 5.0.000 onwards on FLM-0130)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1>?	(f/w ver. 6.0.000 onwards on FLM-0141)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1>?	(f/w ver. 7.0.000 onwards on FLM-0151)

:SYSTem:PCONtrol:MODE<0|1>:SPOint<0|1>?	(f/w ver. 8.0.000 onwards on FLM-0177)

Parameters:	none

Response:	returns the set-point in bits (Mode 0 & Mode 1)	(f/w version 1.0.xxx series only)

returns the set-point in mW (Mode 0) or mA (Mode 1)	(f/w version 1.1.000 onwards)

#### :RANGe<0[|1|2|3]> Commands

This group of commands is used to set up and query the settings for each range. The range is specified as a suffix to the command header. Note that there should be no space between the header and the range number.

###### :CMAX[0|1] <i>

This command is used to specify the maximum current for the given range. On the twin-pump boards, a separate maximum current setting is provided for each pump.

Syntax:	:SYSTem:PCONtrol:RANGe<0|1|2|3>:CMAX <i>	FLM-0100/3	(f/w version 1.0.001 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX <i>	FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX <i>	FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1> <i> 	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX <i>	FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1> <i>	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1> <i>	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1> <i>	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1> <i>	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	<i>	maximum current in mA

Configured:	yes

Response:	none

###### :CMAX[0|1]?

This command is used to query the maximum current for the given range.

Syntax:	:SYSTem:PCONtrol:RANGe<0|1|2|3>:CMAX?	FLM-0100/3	(f/w version 1.0.001 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX?	FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX?	FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1>?	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX?	FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1>?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1>?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1>?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PCONtrol:RANGe<0>:CMAX<0|1>?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the maximum current in mA

### :PDRive Commands

This group of commands is used to set up and query the piezo drive settings. Note that they are only available on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards.

###### :SFACtor <f>

This command is used to specify the scale factor.

Syntax:	:SYSTem:PDRive:SFACtor <f>

Parameters:	<f>	scale factor

Configured:	yes

Response:	none

###### :SFACtor?

This command is used to query the scale factor.

Syntax:	:SYSTem:PDRive:SFACtor?

Parameters:	none

Response:	returns the scale factor

###### :VOLTage <f>

This command is used to specify the DC voltage applied to the piezo.

Syntax:	:SYSTem:PDRive:VOLTage <f>

Parameters:	<f>	voltage in V

Configured:	yes

Response:	none

###### :VOLTage?

This command is used to query the DC voltage applied to the piezo.

Syntax:	:SYSTem:PDRive:VOLTage?

Parameters:	none

Response:	returns the voltage in V

#### :MODulation Commands

This group of commands is used to set up and query the modulation settings. Note that they are only available on the FLM-0121 and FLM-0125 boards.

###### :AMPLitude <f>

This command is used to specify the peak-to-peak amplitude of the modulation.

Syntax:	:SYSTem:PDRive:MODulation:AMPLitude <f>

Parameters:	<f>	amplitude in V

Configured:	yes

Response:	none

###### :AMPLitude?

This command is used to query the peak-to-peak amplitude of the modulation.

Syntax:	:SYSTem:PDRive:MODulation:AMPLitude?

Parameters:	none

Response:	returns the amplitude in V

###### :FREQuency <f>

This command is used to specify the frequency of the modulation.

Syntax:	:SYSTem:PDRive:MODulation:FREQuency <f>

Parameters:	<f>	frequency in Hz

Configured:	yes

Response:	none

###### :FREQuency?

This command is used to query the frequency of the modulation.

Syntax:	:SYSTem:PDRive:MODulation:FREQuency?

Parameters:	none

Response:	returns the frequency in Hz

###### :STATe <b>

This command is used to specify the modulation state.

Syntax:	:SYSTem:PDRive:MODulation:STATe <b>

Parameters:	<b>	modulation state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the modulation state.

Syntax:	:SYSTem:PDRive:MODulation:STATe?

Parameters:	none

Response:	returns the modulation state	0 = disabled; 1 = enabled

### :PMONitor Commands

This group of commands is used to set up and query the pump monitor settings.

###### :C[0|1]0 <f>

###### :C[0|1]1 <f>

These commands are used to specify current coefficients C0 and C1 on the single-pump boards, or coefficients C00, C01, C10 and C11 on the twin-pump boards.

Syntax:	:SYSTem:PMONitor:C0 <f>		FLM-0100	(f/w version 1.0.001 onwards)

:SYSTem:PMONitor:C1 <f>		FLM-0100	(f/w version 1.0.001 onwards)

:SYSTem:PMONitor:C<0|1>0 <f>	FLM-0103	(f/w version 2.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1 <f>	FLM-0103	(f/w version 2.0.000 onwards)

:SYSTem:PMONitor:C0 <f>		FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PMONitor:C1 <f>		FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PMONitor:C0 <f>		FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C1 <f>		FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0 <f>	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1 <f>	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C0 <f>		FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C1 <f>		FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0 <f>	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1 <f>	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0 <f>	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1 <f>	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0 <f>	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1 <f>	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0 <f>	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1 <f>	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	<f>	current coefficient C[0|1]0 (mA) or C[0|1]1 (mA/bit), as appropriate

Configured:	yes

Response:	none

###### :C[0|1]0?

###### :C[0|1]1?

These commands are used to query current coefficients C0 and C1 on the single-pump boards, or coefficients C00, C01, C10 and C11 on the twin-pump boards.

Syntax:	:SYSTem:PMONitor:C0?		FLM-0100	(f/w version 1.0.001 onwards)

:SYSTem:PMONitor:C1?		FLM-0100	(f/w version 1.0.001 onwards)

:SYSTem:PMONitor:C<0|1>0?	FLM-0103	(f/w version 2.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1?	FLM-0103	(f/w version 2.0.000 onwards)

:SYSTem:PMONitor:C0?		FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PMONitor:C1?		FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PMONitor:C0?		FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C1?		FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0?	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1?	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:C0?		FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C1?		FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:C<0|1>0?	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:C<0|1>1?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns temperature coefficient C[0|1]0 (mA) or C[0|1]1 (mA/bit), as appropriate

###### :CURRent[0|1]?

This command is used to query the monitored current. On the twin-pump boards, a separate current monitor is provided for each pump.

Syntax:	:SYSTem:PMONitor:CURRent?	FLM-0100	(f/w version 1.0.001 onwards)

:SYSTem:PMONitor:CURRent<0|1>?	FLM-0103	(f/w version 2.0.000 onwards)

:SYSTem:PMONitor:CURRent?	FLM-0105	(f/w version 3.0.000 onwards)

:SYSTem:PMONitor:CURRent?	FLM-0106	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:CURRent<0|1>?	FLM-0107	(f/w version 4.0.000 onwards)

:SYSTem:PMONitor:CURRent?	FLM-0121/5	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:CURRent<0|1>?	FLM-0130	(f/w version 5.0.000 onwards)

:SYSTem:PMONitor:CURRent<0|1>?	FLM-0141	(f/w version 6.0.000 onwards)

:SYSTem:PMONitor:CURRent<0|1>?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:CURRent<0|1>?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the current in mA

###### :P[0|1]0 <f>

###### :P[0|1]1 <f>

These commands are used to specify power coefficients P[0|1]0 and P[0|1]1. Note that the FLM-0151 and FLM-0177 boards can be fitted with a second power monitor, allowing the power to be monitored before (0) and after (1) the variable optical attenuator.

Syntax:	:SYSTem:PMONitor:P<0|1>0 <f>	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:P<0|1>1 <f>	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:P<0|1>0 <f>	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:P<0|1>1 <f>	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:P0 <f>		others	(f/w version 1.0.001 onwards)

:SYSTem:PMONitor:P1 <f>		others	(f/w version 1.0.001 onwards)

Parameters:	<f>	power coefficient P[0|1]0 (mW) or P[0|1]1 (mW/bit), as appropriate

Configured:	yes

Response:	none

###### :P[0|1]0?

###### :P[0|1]1?

These commands are used to query power coefficients P[0|1]0 and P[0|1]1.

Syntax:	:SYSTem:PMONitor:P<0|1>0?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:P<0|1>1?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:P<0|1>0?	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:P<0|1>1?	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:P0?		others	(f/w version 1.0.001 onwards)

:SYSTem:PMONitor:P1?		others	(f/w version 1.0.001 onwards)

Parameters:	none

Response:	returns power coefficient P[0|1]0 (mW) or P[0|1]1 (mW/bit), as appropriate

###### :POWer[0|1]?

This command is used to query the monitored power at the fiber laser output. Note that the FLM-0151 and FLM-0177 boards can be fitted with a second power monitor, allowing the power to be monitored before (0) and after (1) the variable optical attenuator.

Syntax:	:SYSTem:PMONitor:POWer<[0]|1>?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:POWer<[0]|1>?	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:POWer?		others	(f/w version 1.0.001 onwards)

Parameters:	none

Response:	returns the power in mW

###### :TFACtor[0|1] <f>

This command is used to specify the tap factor(s) associated with the output power monitor(s). The tap factor is the reciprocal of the tap ratio. For example, a 1% tap would require a tap factor of 100, while a 5% tap would require a tap factor of 20. Note that the FLM-0151 and FLM-0177 boards can be fitted with a second power monitor, allowing the power to be monitored before (0) and after (1) the variable optical attenuator.

Syntax:	:SYSTem:PMONitor:TFACtor<0|1> <f>	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:TFACtor<0|1> <f>	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:TFACtor <f>	others	(f/w version 1.0.004 onwards)

Parameters:	<f>	tap factor

Configured:	yes

Response:	none

###### :TFACtor?

This command is used to query the output tap factor(s) associated with the output power monitor(s).

Syntax:	:SYSTem:PMONitor:TFACtor<[0]|1>?	FLM-0151	(f/w version 7.0.000 onwards)

:SYSTem:PMONitor:TFACtor<[0]|1>?	FLM-0177	(f/w version 8.0.000 onwards)

:SYSTem:PMONitor:TFACtor?	others	(f/w version 1.0.004 onwards)

Parameters:	none

Response:	returns the tap factor

### :RIN Commands

This group of commands is used to set up and query the RIN suppression settings. They are not available on the FLM-0105 ASE Controller.

###### :APASs <i>

This command is used to specify the all pass network setting. It is available on the FLM-0141 board only.

Syntax:	:SYSTem:RIN:APASs <i>

Parameters:	<i>	all pass network setting

Configured:	yes

Response:	none

###### :APASs?

This command is used to query the all pass network setting. It is available on the FLM-0141 board only.

Syntax:	:SYSTem:RIN:APASs?

Parameters:	none

Response:	returns the all pass network setting

###### :FLAGs?

This command is used to query the RIN suppression status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:RIN:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.15

Table 5.15 – RIN Suppression Flags

| Bit | Description |
| --- | --- |
| 0 | Mode (0 = Off; 1 = On) |

Note:	All bits are active high.

###### :GAIN <i>

This command is used to specify the gain setting. It is only available on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards.

Syntax:	:SYSTem:RIN:GAIN <i>

Parameters:	<i>	gain setting

Configured:	yes

Response:	none

###### :GAIN?

This command is used to query the gain setting. It is only available on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards.

Syntax:	:SYSTem:RIN:GAIN?

Parameters:	none

Response:	returns the gain setting

###### :HPASs <i>

This command is used to specify the high pass filter setting. It is available on the FLM-0141 board only.

Syntax:	:SYSTem:RIN:HPASs <i>

Parameters:	<i>	high pass filter setting

Configured:	yes

Response:	none

###### :HPASs?

This command is used to query the high pass filter setting. It is available on the FLM-0141 board only.

Syntax:	:SYSTem:RIN:HPASs?

Parameters:	none

Response:	returns the high pass filter setting

###### :MSELect <i>

This command is used to specify the suppression mode.

Syntax:	:SYSTem:RIN:MSELect <i>

Parameters:	<i>	mode select code	see Table 5.16

Configured:	yes

Response:	none

Table 5.16 – RIN Suppression Modes

| Code | Description |
| --- | --- |
| 0 | Off |
| 1 | On |
| 2 | Hardware select (via jumper) |

Note:	Mode 2 unavailable on FLM-0121, FLM-0125 & FLM-0141.

###### :MSELect?

This command is used to query the suppression mode.

Syntax:	:SYSTem:RIN:MSELect?

Parameters:	none

Response:	returns the mode select code 	see Table 5.16

#### :ADJust Commands

This group of commands is used to set up and query the automatic adjustment settings. These commands apply to the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards only.

###### :FLAGs?

This command is used to query the automatic adjustment status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:RIN:ADJust:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.18

Table 5.17 – RIN Auto-Adjust Flags

| Bit | Description |
| --- | --- |
| 0 | RIN stable |
| 1 | Gain normal |
| 2 | Temperature change |

Note:	All bits are active high.

###### :GAIN?

This command is used to query the gain setting currently being used for RIN suppression.

Syntax:	:SYSTem:RIN:ADJust:GAIN?

Parameters:	none

Response:	returns the gain

###### :MONitor?

This command is used to query the internal RIN monitor level.

Syntax:	:SYSTem:RIN:ADJust:MONitor?

Parameters:	none

Response:	returns the RIN monitor level

###### :STATe <b>

This command is used to specify the automatic adjustment state.

Syntax:	:SYSTem:RIN:ADJust:STATe <b>

Parameters:	<b>	auto-adjust state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the automatic adjustment state.

Syntax:	:SYSTem:RIN:ADJust:STATe?

Parameters:	none

Response:	returns the auto-adjust state	0 = disabled; 1 = enabled

### :RTIMe<0|1> Commands

This group of commands is used to set up and query the run time counters. They are only available on the FLM-0141 (Rev. D onwards), FLM-0151 and FLM-0177 boards.

###### :RTIMe<0> <i>

This command is used to clear the total run time counter. It uses a special password protection mechanism and is reserved for factory use only.

Syntax:	:SYSTem:RTIMe<0> <i>

Parameters:	<i>	0 to clear counter

Configured:	yes

Response:	none

###### :RTIMe<0|1>?

This command is used to query the run time counters.

Syntax:	:SYSTem:RTIMe<0>?	Total run time

:SYSTem:RTIMe<1>?	Time since last optimization

Parameters:	none

Response:	returns the run time in hours

### :SMODe Commands

This group of commands is used to set up and query the second mode settings. Note that they are not available on the FLM-0105 ASE Controller.

###### :FLAGs?

This command is used to query the second mode status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:SMODe:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.18

Table 5.18 – Second Mode Flags

| Bit | Description |
| --- | --- |
| 0 | Second mode detected |

Note:	All bits are active high.

###### :LEDDisable <b>

This command is used to specify the LED disable state.

Syntax:	:SYSTem:SMODe:LEDDisable <b>

Parameters:	<b>	LED disable state	1 = LED disabled; 0 = LED enabled

Configured:	yes

Response:	none

###### :LEDDisable?

This command is used to query the LED disable state.

Syntax:	:SYSTem:SMODe:LEDDisable?

Parameters:	none

Response:	returns the LED disable state	1 = LED disabled; 0 = LED enabled

###### :LEVel?

This command is used to query the second mode level.

Syntax:	:SYSTem:SMODe:LEVel?

Parameters:	none

Response:	returns the second mode level in bits

###### :THReshold <i>

This command is used to specify the second mode threshold. When the second mode level is at or above this level, the associated flag is set and the second mode LED is illuminated.

Syntax:	:SYSTem:SMODe:THReshold <i>

Parameters:	<i>	second mode threshold in bits

Configured:	yes

Response:	none

###### :THReshold?

This command is used to query the second mode threshold.

Syntax:	:SYSTem:SMODe:THReshold?

Parameters:	none

Response:	returns the second mode threshold in bits

#### :CORRection Commands

This group of commands is used to set up and query the second mode correction settings.

###### :FLAGs?

This command is used to query the second mode correction status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:SMODe:CORRection:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.18

Table 5.19 – Second Mode Correction Flags

| Bit | Description |
| --- | --- |
| 0 | Positive step applied |
| 1 | Negative step applied |
| 2 | Opposite step applied |
| 3 | Correction stopped to avoid loop |
| 4 | Correction suspended |

Note:	All bits are active high.

###### :HOLDoff <i>

This command is used to specify the hold-off time used by the correction algorithm. The command applies to the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards only.

Syntax:	:SYSTem:SMODe:CORRection:HOLDoff <i>

Parameters:	<i>	hold-off time in seconds

Configured:	yes

Response:	none

###### :HOLDoff?

This command is used to query the hold-off time used by the correction algorithm. It applies to the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards only.

Syntax:	:SYSTem:SMODe:CORRection:HOLDoff?

Parameters:	none

Response:	returns the hold-off time in seconds

###### :RANGe <i>

This command is used to specify the range over which the temperature of the wide-band grating may be adjusted. The value represents the total (peak-to-peak) range and is expressed in terms of the step size. The command applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:SMODe:CORRection:RANGe <i>

Parameters:	<i>	range in steps

Configured:	yes

Response:	none

###### :RANGe?

This command is used to query the range over which the temperature of the wide-band grating may be adjusted. It applies to the FLM-0100, FLM-0103, FLM-0106 and FLM-0107 boards only.

Syntax:	:SYSTem:SMODe:CORRection:RANGe?

Parameters:	none

Response:	returns the range in steps

###### :STATe <b>

This command is used to specify the correction state.

Syntax:	:SYSTem:SMODe:CORRection:STATe <b>	(f/w version 4.2.000 to 6.2.001)

:SYSTem:SMODe:CORRection:STATe <i>	(f/w version 6.3.000 onwards)

Parameters:	<b>	correction state	see Table 5.20	(f/w version 4.2.000 to 6.2.001)

<i>	correction state	see Table 5.20	(f/w version 6.3.000 onwards)

Configured:	yes

Response:	none

Table 5.20 – Second Mode Correction States

| State | Correction State | Firmware Version |
| --- | --- | --- |
| 0 | Disabled | 4.2.000 onwards |
| 1 | Enabled – step correction method | 4.2.000 onwards |
| 2 | Enabled – optimization method | 6.3.000 onwards |

###### :STATe?

This command is used to query the correction state.

Syntax:	:SYSTem:SMODe:CORRection:STATe?

Parameters:	none

Response:	returns the correction state	see Table 5.20

###### :STEP <f>

This command is used to specify the correction step size.

Syntax:	:SYSTem:SMODe:CORRection:STEP <f>

Parameters:	<f>	step size in ºC		FLM-0101, FLM-0103, FLM-0106 & FLM-0107

<f>	step size in V		FLM-0121, FLM-0125, FLM-0130, FLM-0141,

FLM-0151 & FLM-0177

Configured:	yes

Response:	none

###### :STEP?

This command is used to query the correction step size.

Syntax:	:SYSTem:SMODe:CORRection:STEP?

Parameters:	none

Response:	returns the step size in ºC		FLM-0101, FLM-0103, FLM-0106 & FLM-0107

returns the step size in V		FLM-0121, FLM-0125, FLM-0130, FLM-0141,

FLM-0151 & FLM-0177

##### :RAMP Commands

This group of commands is used to set up and query the ramp settings used for the second mode correction process. The ramp comprises a fixed pattern of 20 steps and is applied to the piezo to determine the correction required to move away from the second mode region. Although the ramp pattern is fixed, its overall amplitude and duration can be adjusted. Note that these commands are only available on the FLM-0121, FLM-0125, FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards.

###### :AMPLitude <f>

This command is used to specify the peak amplitude of the ramp.

Syntax:	:SYSTem:SMODe:CORRection:RAMP:AMPLitude <f>

Parameters:	<f>	amplitude in V

Configured:	yes

Response:	none

###### :AMPLitude?

This command is used to query the peak amplitude of the ramp.

Syntax:	:SYSTem:SMODe:CORRection:RAMP:AMPLitude?

Parameters:	none

Response:	returns the amplitude in V

###### :DWELl <f>

This command is used to specify the dwell time at each step of the ramp.

Syntax:	:SYSTem:SMODe:CORRection:RAMP:DWELl <i>

Parameters:	<i>	dwell time in 1/10ths of a second

Configured:	yes

Response:	none

###### :DWELl?

This command is used to query the dwell time at each step of the ramp.

Syntax:	:SYSTem:SMODe:CORRection:RAMP:DWELl?

Parameters:	none

Response:	returns the dwell time in 1/10ths of a second

### :TCONtrol<0[|1|2|3|4]> Commands

This group of commands is used to set up and query the temperature control settings. The control channel is specified as a suffix to the command header. Note that there should be no space between the header and the channel number. The valid channel numbers are defined in Table 5.21.

Table 5.21 – Temperature Control Channels

| Channel | FLM-0100 (single-pump) | FLM-0103 (twin-pump) | FLM-0105 (ASE) | FLM-0106 (single-pump) | FLM-0107 (twin-pump) | FLM-0121/5 (single-pump) | FLM-0130 (twin-pump) | FLM-0141 (twin-pump) | FLM-0151 (twin-pump) | FLM-0177 (twin-pump) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | TEC | TEC 1 | TEC | TEC | TEC 1 | TEC | TEC 1 | TEC 1 | TEC 1 | TEC 1 |
| 1 | Heater 1 | TEC 2 | - | Heater 1 | TEC 2 | Heater | TEC 2 | TEC 2 | TEC 2 | TEC 2 |
| 2 | Heater 2 | Heater 1 | - | Heater 2 | Heater 1 | PCB | Heater | Heater | Heater | Heater |
| 3 | PCB | Heater 2 | - | PCB | Heater 2 | - | PCB | PCB | PCB | PCB |
| 4 | - | PCB | - | - | PCB | - | - | - | - | see text |

Note:	FLM-0100 channel assignments require firmware version 1.0.000 onwards.

FLM-0103 channel assignments require firmware version 2.0.000 onwards.

FLM-0105 channel assignments require firmware version 3.0.000 onwards.

FLM-0106 channel assignments require firmware version 4.0.000 onwards.

FLM-0107 channel assignments require firmware version 4.0.000 onwards.

FLM-0121/5 channel assignments require firmware version 5.0.000 onwards.

FLM-0130 channel assignments require firmware version 5.0.000 onwards.

FLM-0141 channel assignments require firmware version 6.0.000 onwards.

FLM-0151 channel assignments require firmware version 7.0.000 onwards.

FLM-0177 channel assignments require firmware version 8.0.000 onwards.

PCB temperature references require firmware version 4.7.000 onwards.

In the case of the PCB temperature channel, only a subset of this group of commands is available. Specifically, these are the :FLAGs?, :TMAX <f> and :TMAX? commands.

In the case of the FLM-0177, when configured for the Power Rock application, Channel 1 is used for the amplifier TEC (TEC 2). Additionally, Channel 4 is used for the amplifier PCB (PCB 2).

###### :DELay <i>

This command is used to specify the heater startup delay. It applies to the FLM-0141 board only and allows the startup of individual channels on multi-laser systems to be staggered, so as to even out power surges during warm-up. Note that the actual delay will depend on the position of the laser module on its line card. For example, if the delay is specified as 60 seconds on all modules, then the module at address 0 will start up immediately after power-up or reset, the module at address 1 will start up after 60 seconds, the module at address 2 will start up after 120 seconds, and so on.

Syntax:	:SYSTem:TCONtrol<2>:DELay <i>

Parameters:	<i>	startup delay in s

Configured:	yes

Response:	none

###### :DELay?

This command is used to query the heater startup delay.

Syntax:	:SYSTem:TCONtrol<2>:DELay?

Parameters:	none

Response:	returns the startup delay in s

###### :DOUTput <i>

This command is used to specify the heater output value directly and is intended for test purposes. It applies to the FLM-0130 (Rev. B onwards), FLM-0141, FLM-0151 and FLM-0177 boards only.

Syntax:	:SYSTem:TCONtrol<2>:DOUTput <i>

Parameters:	<i>	output value	0 – 4095 (direct output); -1 (normal operation)

Configured:	no

Response:	none

###### :DOUTput?

This command is used to query the heater output value. It applies to the FLM-0130 (Rev. B onwards), FLM-0141, FLM-0151 and FLM-0177 boards only.

Syntax:	:SYSTem:TCONtrol<2>:DOUTput?

Parameters:	none

Response:	returns the output value

###### :ELEMents<0|1> <i>

This command is used to specify the heater element configurations. It applies to the FLM-0130 (Rev. B onwards), FLM-0151 and FLM-0177 boards only.

Syntax:	:SYSTem:TCONtrol<2>:ELEMents<0> <i>	normal mode 	(temperature error  <  1°C)

:SYSTem:TCONtrol<2>:ELEMents<1> <i>	boost mode	(temperature error >= 1°C)

Parameters:	<i>	element configuration	see Table 5.22

Configured:	yes

Response:	none

###### :ELEMents<0|1>?

This command is used to query the heater element configurations. It applies to the FLM-0130 (Rev. B onwards), FLM-0151 and FLM-0177 boards only.

Syntax:	:SYSTem:TCONtrol<2>:ELEMents<0>?	normal mode 	(temperature error  <  1°C)

:SYSTem:TCONtrol<2>:ELEMents<1>?	boost mode	(temperature error >= 1°C)

Parameters:	none

Response:	returns the element configuration

Table 5.22 – Heater Element Configuration

| Code | Element 1 | Element 2 | Element 3 |
| --- | --- | --- | --- |
| 0 | disabled | disabled | disabled |
| 1 | enabled | disabled | disabled |
| 2 | disabled | enabled | disabled |
| 3 | enabled | enabled | disabled |
| 4 | disabled | disabled | enabled |
| 5 | enabled | disabled | enabled |
| 6 | disabled | enabled | enabled |
| 7 | enabled | enabled | enabled |

###### :FLAGs?

This command is used to query the temperature control status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3|4]>:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.23

Table 5.23 – Temperature Control Flags

| Bit | Description |
| --- | --- |
| 0 | Temperature stable |
| 1 | Alarm 1 (TEC only on FLM-0100/5/6; all channels on FLM-0103/7) |
| 2 | Alarm 2 (TEC only on FLM-0100/5/6; all channels on FLM-0103/7) |
| 3 | Over-temperature |
| 4 | Under-temperature |
| 5 | Shutdown |

Note:	All bits are active high.

###### :KD <f>

This command is used to specify the differential constant used by the control loop.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:KD <f>

Parameters:	<f>	differential constant in bits/ºC.s

Configured:	yes

Response:	none

###### :KD?

This command is used to query the differential constant used by the control loop.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:KD?

Parameters:	none

Response:	returns the differential constant in bits/ºC.s

###### :KI <f>

This command is used to specify the integral constant used by the control loop.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:KI <f>

Parameters:	<f>	integral constant in bits/ºC/s

Configured:	yes

Response:	none

###### :KI?

This command is used to query the integral constant used by the control loop.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:KI?

Parameters:	none

Response:	returns the integral constant in bits/ºC/s

###### :KP <f>

This command is used to specify the proportional constant used by the control loop.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:KP <f>

Parameters:	<f>	proportional constant in bits/ºC

Configured:	yes

Response:	none

###### :KP?

This command is used to query the proportional constant used by the control loop.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:KP?

Parameters:	none

Response:	returns the proportional constant in bits/ºC

###### :OMAX <f>

This command is used to specify the maximum output value used by the control loop. It applies to the Power Rock only and is used to limit the amplifier’s TEC drive current.

Syntax:	:SYSTem:TCONtrol1:OMAX <f>	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:TCONtrol2:OMAX <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	maximum output value in bits

Configured:	yes

Response:	none

###### :OMAX?

This command is used to query the maximum output value used by the control loop. It applies to the Power Rock only and is used to limit the amplifier’s TEC drive current.

Syntax:	:SYSTem:TCONtrol1:OMAX?		Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:TCONtrol2:OMAX?		Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns the maximum output value in bits

###### :OMIN <f>

This command is used to specify the minimum output value used by the control loop. It applies to the Power Rock only and is used to limit the amplifier’s TEC drive current.

Syntax:	:SYSTem:TCONtrol1:OMIN <f>	Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:TCONtrol2:OMIN <f>	Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	<f>	minimum output value in bits

Configured:	yes

Response:	none

###### :OMIN?

This command is used to query the minimum output value used by the control loop. It applies to the Power Rock only and is used to limit the amplifier’s TEC drive current.

Syntax:	:SYSTem:TCONtrol1:OMIN?		Power Rock	(f/w version 8.1.000 onwards)

:SYSTem:TCONtrol2:OMIN?		Amp. ASE	(f/w version 8.5.000 onwards)

Parameters:	none

Response:	returns the minimum output value in bits

###### :PWMPeriod <i>

This command is used to specify the heater PWM period. It applies to the FLM-0141 board only.

Syntax:	:SYSTem:TCONtrol<2>:PWMPeriod <i>

Parameters:	<i>	PWM period

Configured:	yes

Response:	none

###### :PWMPeriod?

This command is used to query the heater PWM period. It applies to the FLM-0141 board only.

Syntax:	:SYSTem:TCONtrol<2>:PWMPeriod?

Parameters:	none

Response:	returns the PWM period

###### :SPOint <f>

This command is used to specify the temperature set-point.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:SPOint <f>

Parameters:	<f>	set-point in ºC

Configured:	yes

Response:	none

###### :SPOint?

This command is used to query the temperature set-point.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:SPOint?

Parameters:	none

Response:	returns the set-point in ºC

###### :STATe <b>

This command is used to specify the control loop state.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:STATe <b>

Parameters:	<b>	control loop state	0 = disabled; 1 = enabled

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the control loop state.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:STATe?

Parameters:	none

Response:	returns the control loop state	0 = disabled; 1 = enabled

###### :TMAX <f>

This command is used to specify the over-temperature alarm limit. If the monitored temperature exceeds this value, then the over-temperature and shutdown flags are set and the control loop is shut down. In the case of the PCB over-temperature alarm limit being exceeded, all of the control loops are shut down.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3|4]>:TMAX <f>

Parameters:	<f>	over-temperature limit in ºC

Configured:	yes

Response:	none

###### :TMAX?

This command is used to query the over-temperature alarm limit.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3|4]>:TMAX?

Parameters:	none

Response:	returns the over-temperature limit in ºC

###### :TMIN <f>

This command is used to specify the under-temperature alarm limit. If the monitored temperature falls below this value, then the under-temperature and shutdown flags are set and the control loop is shut down.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:TMIN <f>

Parameters:	<f>	under-temperature limit in ºC

Configured:	yes

Response:	none

###### :TMIN?

This command is used to query the under -temperature alarm limit.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:TMIN?

Parameters:	none

Response:	returns the under -temperature limit in ºC

###### :TOLerance <f>

This command is used to specify the temperature tolerance. When the difference between the measured temperature and the set-point is within this limit, the temperature is considered to be stable. When all enabled temperature control loops are stable, the Temperature Stable LED is illuminated.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:TOLerance <f>

Parameters:	<f>	tolerance in ºC

Configured:	yes

Response:	none

###### :TOLerance?

This command is used to query the temperature tolerance.

Syntax:	:SYSTem:TCONtrol<0[|1|2|3]>:TOLerance?

Parameters:	none

Response:	returns the tolerance in ºC

###### :VMONitor[1|2|3]?

This command is used to query the heater element voltages. It applies to the FLM-0130 (Rev. B onwards) and FLM-0151 boards only. On the FLM-0130, the voltages can be monitored individually; on the FLM-0151, they are summed together.

Syntax:	:SYSTem:TCONtrol<2>:VMONitor<1|2|3>?	FLM-0130	(f/w version 5.1.000 onwards)

:SYSTem:TCONtrol<2>:VMONitor?	FLM-0151	(f/w version 7.0.000 onwards)

Parameters:	none

Response:	returns the voltage in V

### :TMONitor<0[|1|2|3|4]> Commands

This group of commands is used to set up and query the temperature monitor settings. The monitor channel is specified as a suffix to the command header. Note that there should be no space between the header and the channel number. The valid channel numbers are defined in Table 5.24.

Table 5.24 – Temperature Monitor Channels

| Channel | FLM-0100 (single-pump) | FLM-0103 (twin-pump) | FLM-0105 (ASE) | FLM-0106 (single-pump) | FLM-0107 (twin-pump) | FLM-0121/5 (single-pump) | FLM-0130 (twin-pump) | FLM-0141 (twin-pump) | FLM-0151 (twin-pump) | FLM-0177 (twin-pump) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 0 | TEC | TEC 1 | TEC | TEC | TEC 1 | TEC | TEC 1 | TEC 1 | TEC 1 | TEC 1 |
| 1 | Heater 1 | TEC 2 | - | Heater 1 | TEC 2 | Heater | TEC 2 | TEC 2 | TEC 2 | TEC 2 |
| 2 | Heater 2 | Heater 1 | - | Heater 2 | Heater 1 | PCB | Heater | Heater | Heater | Heater |
| 3 | PCB | Heater 2 | - | PCB | Heater 2 | - | PCB | PCB | PCB | PCB |
| 4 | - | PCB | - | - | PCB | - | - | - | - | - |

Note:	FLM-0100 channel assignments require firmware version 1.0.000 onwards.

FLM-0103 channel assignments require firmware version 2.0.000 onwards.

FLM-0105 channel assignments require firmware version 3.0.000 onwards.

FLM-0106 channel assignments require firmware version 4.0.000 onwards.

FLM-0107 channel assignments require firmware version 4.0.000 onwards.

FLM-0121/5 channel assignments require firmware version 5.0.000 onwards.

FLM-0130 channel assignments require firmware version 5.0.000 onwards.

FLM-0141 channel assignments require firmware version 6.0.000 onwards.

FLM-0151 channel assignments require firmware version 7.0.000 onwards.

FLM-0177 channel assignments require firmware version 8.0.000 onwards.

###### :FLAGs?

This command is used to query the temperature monitor status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:TMONitor<0[|1|2|3|4]>:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.25

Table 5.25 – Temperature Monitor Flags

| Bit | Description |
| --- | --- |
| 0 | Sensor fault |

Note:	All bits are active high.

###### :T0 <f>

###### :T1 <f>

###### :T2 <f>

###### :T3 <f>

These commands are used to specify temperature coefficients T0, T1, T2, and T3.

Syntax:	:SYSTem:TMONitor<0[|1|2|3|4]>:T0 <f>

:SYSTem:TMONitor<0[|1|2|3|4]>:T1 <f>

:SYSTem:TMONitor<0[|1|2|3|4]>:T2 <f>

:SYSTem:TMONitor<0[|1|2|3|4]>:T3 <f>

Parameters:	<f>	temperature coefficient T0 (ºC), T1 (ºC/bit), T2 (ºC/bit2), or T3 (ºC/bit3), as appropriate

Configured:	yes

Response:	none

###### :T0?

###### :T1?

###### :T2?

###### :T3?

These commands are used to query temperature coefficients T0, T1, T2, and T3.

Syntax:	:SYSTem:TMONitor<0[|1|2|3|4]>:T0?

:SYSTem:TMONitor<0[|1|2|3|4]>:T1?

:SYSTem:TMONitor<0[|1|2|3|4]>:T2?

:SYSTem:TMONitor<0[|1|2|3|4]>:T3?

Parameters:	none

Response:	returns temperature coefficient T0 (ºC), T1 (ºC/bit), T2 (ºC/bit2), or T3 (ºC/bit3), as appropriate

###### :TEMPerature?

This command is used to query the monitored temperature.

Syntax:	:SYSTem:TMONitor<0[|1|2|3|4]>:TEMPerature?

Parameters:	none

Response:	returns the temperature in ºC

### :VOA Commands

This group of commands is used to set up and query the variable optical attenuator (VOA) settings. Note that they are only available on the FLM-0151 and FLM-0177 boards.

###### :ATTenuation <f>

This command is used to specify the attenuation. The value is specified as a multiplying factor in the range 0 (maximum attenuation) to 1 (minimum attenuation).

Syntax:	:SYSTem:VOA:ATTenuation <f>

Parameters:	<f>	attenuation

Configured:	yes

Response:	none

###### :ATTenuation?

This command is used to query the attenuation.

Syntax:	:SYSTem:VOA:ATTenuation?

Parameters:	none

Response:	returns the attenuation

###### :VOLTage <f>

This command is used to specify the DC voltage applied to the VOA.

Syntax:	:SYSTem:VOA:VOLTage <f>

Parameters:	<f>	voltage in V

Configured:	yes

Response:	none

###### :VOLTage?

This command is used to query the DC voltage applied to the VOA.

Syntax:	:SYSTem:VOA:VOLTage?

Parameters:	none

Response:	returns the voltage in V

#### :DATa<i> Commands

This group of commands is used to set up and query the calibration data. The data point number, which should be in the range 0 to 20, is specified as a suffix to the command header. Note that there should be no space between the header and the suffix.

###### :ATTenuation <f>

This command is used to specify the attenuation at the data point of interest. The value is specified as a multiplying factor in the range 0 (maximum attenuation) to 1 (minimum attenuation).

Syntax:	:SYSTem:VOA:DATa<i>:ATTenuation <f>

Parameters:	<f>	attenuation

Configured:	yes

Response:	none

###### :ATTenuation?

This command is used to query the attenuation at the data point of interest.

Syntax:	:SYSTem:VOA:DATa<i>:ATTenuation?

Parameters:	none

Response:	returns the attenuation

###### :VOLTage <f>

This command is used to specify the voltage at the data point of interest.

Syntax:	:SYSTem:VOA:DATa<i>:VOLTage <f>

Parameters:	<f>	voltage in V

Configured:	yes

Response:	none

###### :VOLTage?

This command is used to query the voltage at the data point of interest.

Syntax:	:SYSTem:VOA:DATa<i>:VOLTage?

Parameters:	none

Response:	returns the voltage in V

#### :MODulation Commands

This group of commands is used to set up and query the modulation settings. Note that they are only available on the FLM-0151 board.

###### :AMPLitude <f>

This command is used to specify the peak-to-peak amplitude of the internally generated waveform.

Syntax:	:SYSTem:VOA:MODulation:AMPLitude <f>

Parameters:	<f>	amplitude in V

Configured:	yes

Response:	none

###### :AMPLitude?

This command is used to query the peak-to-peak amplitude of the internally generated waveform.

Syntax:	:SYSTem:VOA:MODulation:AMPLitude?

Parameters:	none

Response:	returns the amplitude in V

###### :FREQuency <f>

This command is used to specify the frequency of the internally generated waveform.

Syntax:	:SYSTem:VOA:MODulation:FREQuency <f>

Parameters:	<f>	frequency in Hz

Configured:	yes

Response:	none

###### :FREQuency?

This command is used to query the frequency of the internally generated waveform.

Syntax:	:SYSTem:VOA:MODulation:FREQuency?

Parameters:	none

Response:	returns the frequency in Hz

###### :STATe <i>

This command is used to specify the modulation state.

Syntax:	:SYSTem:VOA:MODulation:STATe <i>

Parameters:	<i>	modulation state	see Table 5.26

Configured:	yes

Response:	none

Table 5.26 – VOA Modulation States

| State | Description |
| --- | --- |
| 0 | None |
| 1 | Internal (waveform generator) |
| 2 | External |

###### :STATe?

This command is used to query the modulation state.

Syntax:	:SYSTem:VOA:MODulation:STATe?

Parameters:	none

Response:	returns the modulation state	see Table 5.26

###### :WAVeform <i>

This command is used to specify the type of internally generated waveform.

Syntax:	:SYSTem:VOA:MODulation:WAVeform <i>

Parameters:	<i>	waveform	see Table 5.27

Configured:	yes

Response:	none

Table 5.27 – VOA Modulation Waveforms

| Type | Description |
| --- | --- |
| 0 | Sine wave |
| 1 | Triangle wave |
| 2 | Square wave |

###### :WAVeform?

This command is used to query the type of internally generated waveform.

Syntax:	:SYSTem:VOA:MODulation:WAVeform?

Parameters:	none

Response:	returns the waveform	see Table 5.27

### :WADJust Commands

This group of commands is used to set up and query the wavelength adjust settings. They allow the fiber laser wavelength to be adjusted via a potentiometer input, a BNC input or from the command interface. Note that these commands are not available on the FLM-0105 ASE Controller.

###### :BAND <a>

This command is used to specify the mode hop free band.

Syntax:	:SYSTem:WADJust:BAND <a>	(f/w version 1.0.003a only)

Parameters:	<a>	band identifier

Configured:	yes

Response:	none

###### :BAND?

This command is used to query the mode hop free band.

Syntax:	:SYSTem:WADJust:BAND?		(f/w version 1.0.003a only)

Parameters:	none

Response:	returns the band identifier

###### :FLAGs?

This command is used to query the wavelength adjust status. The return value is an integer comprising a series of flags.

Syntax:	:SYSTem:WADJust:FLAGs?

Parameters:	none

Response:	returns the status flags	see Table 5.28

Table 5.28 – Wavelength Adjust Flags

| Bit | Description |
| --- | --- |
| 0 | Correction applied |

Note:	All bits are active high.

###### :INPut <i>

This command is used to specify the analog input to be used for adjusting the wavelength. When the potentiometer input is selected, the firmware compensates to give the control knob a linear response. Note that the command is not available on the FLM-0141 board and not used on the FLM-0151 and FLM-0177 boards, which have a linear potentiometer input.

Syntax:	:SYSTem:WADJust:INPut <i>

Parameters:	<i>	input	0 = BNC socket; 1 = potentiometer

Configured:	yes

Response:	none

###### :INPut?

This command is used to query the analog input used for adjusting the wavelength. It is not available on the FLM-0141 board and not used on the FLM-0151 and FLM-0177 boards.

Syntax:	:SYSTem:WADJust:INPut?

Parameters:	none

Response:	returns the input

###### :STATe <b>

###### :STATe <i>

This command is used to specify the adjustment state. This setting works in conjunction with the pump status, or independently, as shown in Table 5.29. When the adjustment state is set to 1 and the pump is disabled, the potentiometer (or BNC) input is sampled and used to set the wavelength; if the pump is then enabled, the wavelength setting is frozen at the last sample point, so as to avoid any effect due to noise at the potentiometer (or BNC) input. When the adjustment state is set to 2, the potentiometer (or BNC) input is used unconditionally, irrespective of the pump status. When the adjustment state is set to 0, the potentiometer (or BNC) input is disabled. Note that the command is not available on the FLM-0141 board.

Syntax:	:SYSTem:WADJust:STATe <b>	(f/w version 1.0.003a & 2.1.000 - 4.0.002)

:SYSTem:WADJust:STATe <i>	(f/w version 4.1.000 onwards)

Parameters:	<b>	adjustment state	see Table 5.29

<i>	adjustment state	see Table 5.29

Configured:	yes

Response:	none

Table 5.29 – Wavelength Adjustment States

| State | Wavelength Adjustment | Firmware Version |
| --- | --- | --- |
| 0 | Disabled | 1.0.003a and 2.1.000 onwards |
| 1 | Enabled when pump is disabled | 1.0.003a and 2.1.000 onwards |
| 2 | Enabled unconditionally | 4.1.000 onwards |

###### :STATe?

This command is used to query the adjustment state. It is not available on the FLM-0141 board.

Syntax:	:SYSTem:WADJust:STATe?		(f/w version 1.0.003a & 2.1.000 onwards)

Parameters:	none

Response:	returns the adjustment state	see Table 5.29

###### :TOFFset?

This command is used to query the temperature offset being applied to the wide-band grating. The value depends on the potentiometer position and is scaled according to the WB half-range setting for the selected band (see Section 5.1.7.1).

Syntax:	:SYSTem:WADJust:TOFFset?	(f/w version 1.0.003a only)

Parameters:	none

Response:	returns the temperature offset in ºC

###### :WAVelength <f>

This command is used to specify the wavelength when the potentiometer (or BNC) input is inactive (wavelength adjustment disabled or pump enabled).

Syntax:	:SYSTem:WADJust:WAVelength <f>

Parameters:	<f>	wavelength in nm

Configured:	no

Response:	none

###### :WAVelength?

This command is used to query the wavelength. It can be used at any time, irrespective of whether the wavelength is being set via the potentiometer input, the BNC input or the command interface.

Syntax:	:SYSTem:WADJust:WAVelength?

Parameters:	none

Response:	returns the wavelength in nm

### :Shutdown Commands

These commands were added for the GEN VI systems to prevent the seed laser from shutting down immediately. These commands have no abbreviations.

:DwellVal <i>

This command sets the length of the dwell value in seconds.

Syntax:	:System:Shutdown:DwellVal <i>

Parameters:	<i> dwell value in seconds

Configured:	yes

Response:	none

:DwellVal?

The command is used to query the dwell value.

Syntax:	:System:Shutdown:DwellVal?

Parameters:	none

Configured:	no

Response:	returns the dwell value in seconds.

:DState?

The command is used to query the the state of the dwell. It returns a 0 if the dwell is not active and 1 if it is. This command is used by FLCP to countdown the amount of dwell remaining.

Syntax:	:System:Shutdown:DwellVal?

Parameters:	None

Configured:	no

Response:	returns the state of the dwell.

:Dmin <i>

This command is used to set the minimum allowed dwell in seconds.

Syntax:	System:Shutdown:Dmin <i>

Parameters:	value of minimum allowed dwell.

Configured:	yes

Response:	none

:Dmin?

This command is used to query the minimum allowed dwell in seconds.

Syntax:	System:Shutdown:Dmin?

Parameters:	none

Configured:	no

Response:	value of minimum allowed dwell

:Dmax <i>

This command is used to query the minimum allowed dwell in seconds.

Syntax:	System:Shutdown:Dmax <i>

Parameters:	value of maximum allowed dwell.

Configured:	yes

Response:	none

:Dmax?

This command is used to query the minimum allowed dwell in seconds.

Syntax:	System:Shutdown:Dmax?

Parameters:	none

Configured:	no

Response:	value of maximum allowed dwell

## :STATus[0|1|2|3] Commands

These commands are used to query the status of the controller. They allow a complete set of monitor readings and/or flags to be reported in a single message, providing a faster alternative to querying the parameters individually. Depending on firmware version, an optional format specifier can be included as a suffix to the command header. This allows additional information to be reported for some board types. Note that there should be no space between the header and the format specifier.

###### :ALL?

This command is used to request the full status report, comprising both monitor readings and flags. The data is returned as a list of comma-separated values. The number of parameters depends on the board type and is summarised in Table 5.30.

Syntax:	:STATus:ALL?	all board types	(f/w version 5.0.000 onwards)

:STATus0:ALL?	all board types	(f/w version 6.4.001 onwards)

:STATus1:ALL?	FLM-0121/5, FLM-0130 & FLM-0141	(f/w version 6.4.001 onwards)

:STATus2:ALL?	FLM-0141, Rev. D onwards	(f/w version 6.5.000 onwards)

:STATus3:ALL?	FLM-0151	(f/w version 7.0.000 onwards)

:STATus3:ALL?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the full status report	see Table 5.30

Table 5.30 – Parameter List returned by :STATus[0|1|2|3]:ALL?

| Parameter | FLM-0100/6 (single-pump) | FLM-0103/7 (twin-pump) | FLM-0105 (ASE) | FLM-0121/5 (single-pump) | FLM-0130 (twin-pump) | FLM-0141 (twin-pump) | FLM-0151 (twin-pump) | FLM-0177 (twin-pump) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Error code |  |  |  |  |  |  |  |  |
| Pump control flags |  |  |  |  |  |  |  |  |
| RIN on/off state |  |  |  |  |  |  |  |  |
| Power monitor |  |  |  |  |  |  |  |  |
| Second mode level |  |  |  |  |  |  |  |  |
| Second mode flag |  |  |  |  |  |  |  |  |
| Second mode correction flags |  |  |  |  |  |  |  |  |
| Pump 0 current monitor |  |  |  |  |  |  |  |  |
| Pump 0 back facet level | 1 | 2 |  |  |  |  |  |  |
| Pump 0 back facet flag | 1 | 2 |  |  |  |  |  |  |
| Pump 1 current monitor |  |  |  |  |  |  |  |  |
| Pump 1 back facet level |  | 2 |  |  |  |  |  |  |
| Pump 1 back facet flag |  | 2 |  |  |  |  |  |  |
| TEC 0 temperature monitor |  |  |  |  |  |  |  |  |
| TEC 0 temperature monitor flags |  |  |  |  |  |  |  |  |
| TEC 0 temperature control flags |  |  |  |  |  |  |  |  |
| TEC 1 temperature monitor |  |  |  |  |  |  |  |  |
| TEC 1 temperature monitor flags |  |  |  |  |  |  |  |  |
| TEC 1 temperature control flags |  |  |  |  |  |  |  |  |
| Heater 0 temperature monitor |  |  |  |  |  |  |  |  |
| Heater 0 temperature monitor flags |  |  |  |  |  |  |  |  |
| Heater 0 temperature control flags |  |  |  |  |  |  |  |  |
| Heater 1 temperature monitor |  |  |  |  |  |  |  |  |
| Heater 1 temperature monitor flags |  |  |  |  |  |  |  |  |
| Heater 1 temperature control flags |  |  |  |  |  |  |  |  |
| PCB temperature monitor |  |  |  |  |  |  |  |  |
| PCB temperature monitor flags |  |  |  |  |  |  |  |  |
| PCB temperature control flags |  |  |  |  |  |  |  |  |
| Heater 0 temperature set-point |  |  |  |  |  |  |  |  |
| Heater 1 temperature set-point |  |  |  |  |  |  |  |  |
| Piezo drive voltage set-point |  |  |  |  |  |  |  |  |
| Wavelength |  |  |  |  |  |  |  |  |
| Wavelength adjust flags |  |  |  |  |  |  |  |  |
| RIN auto-adjust gain |  |  |  | 3 | 3 | 3 | 3 | 3 |
| RIN auto-adjust flags |  |  |  | 3 | 3 | 3 | 3 | 3 |
| RIN auto-adjust monitor level |  |  |  | 3 | 3 | 3 | 3 | 3 |
| Run Time 0 (total run time) |  |  |  |  |  | 4 | 4 | 4 |
| Run Time 1 (time since optimization) |  |  |  |  |  | 4 | 4 | 4 |
| VOA attenuation set-point |  |  |  |  |  |  | 5 | 5 |
| VOA voltage set-point |  |  |  |  |  |  | 5 | 5 |
| VOA voltage monitor |  |  |  |  |  |  | 5 | 5 |
| VOA power mute state |  |  |  |  |  |  | 5 | 5 |
| VOA output power monitor |  |  |  |  |  |  | 5 | 5 |

1 FLM-0106 only

2 FLM-0107 only

3 :STATus1:ALL?, :STATus2:ALL? and :STATus3:ALL? only

4 :STATus2:ALL? and :STATus3:ALL? only

5 :STATus3:ALL? only

###### :FLAGs?

This command is used to request the short status report, comprising flags only. The data is returned as a list of comma-separated values. The number of parameters depends on the board type and is summarised in Table 5.31.

Syntax:	:STATus:FLAGs?	all board types	(f/w version 5.0.000 onwards)

:STATus0:FLAGs?	all board types	(f/w version 6.4.001 onwards)

:STATus1:FLAGs?	FLM-0121/5, FLM-0130 & FLM-0141	(f/w version 6.4.001 onwards)

:STATus2:FLAGs?	FLM-0141, Rev. D onwards	(f/w version 6.5.000 onwards)

:STATus3:FLAGs?	FLM-0151	(f/w version 7.0.000 onwards)

:STATus3:FLAGs?	FLM-0177	(f/w version 8.0.000 onwards)

Parameters:	none

Response:	returns the short status report	see Table 5.31

Table 5.31 – Parameter List returned by :STATus[0|1|2|3]:FLAGs?

| Parameter | FLM-0100/6 (single-pump) | FLM-0103/7 (twin-pump) | FLM-0105 (ASE) | FLM-0121/5 (single-pump) | FLM-0130 (twin-pump) | FLM-0141 (twin-pump) | FLM-0151 (twin-pump) | FLM-0177 (twin-pump) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Error code |  |  |  |  |  |  |  |  |
| Pump control flags |  |  |  |  |  |  |  |  |
| RIN on/off state |  |  |  |  |  |  |  |  |
| Second mode flag |  |  |  |  |  |  |  |  |
| Second mode correction flags |  |  |  |  |  |  |  |  |
| Pump 0 back facet flag | 1 | 2 |  |  |  |  |  |  |
| Pump 1 back facet flag |  | 2 |  |  |  |  |  |  |
| TEC 0 temperature monitor flags |  |  |  |  |  |  |  |  |
| TEC 0 temperature control flags |  |  |  |  |  |  |  |  |
| TEC 1 temperature monitor flags |  |  |  |  |  |  |  |  |
| TEC 1 temperature control flags |  |  |  |  |  |  |  |  |
| Heater 0 temperature monitor flags |  |  |  |  |  |  |  |  |
| Heater 0 temperature control flags |  |  |  |  |  |  |  |  |
| Heater 1 temperature monitor flags |  |  |  |  |  |  |  |  |
| Heater 1 temperature control flags |  |  |  |  |  |  |  |  |
| PCB temperature monitor flags |  |  |  |  |  |  |  |  |
| PCB temperature control flags |  |  |  |  |  |  |  |  |
| Wavelength adjust flags |  |  |  |  |  |  |  |  |
| RIN auto-adjust flags |  |  |  | 3 | 3 | 3 | 3 | 3 |
| VOA power mute state |  |  |  |  |  |  | 4 | 4 |

1 FLM-0106 only

2 FLM-0107 only

3 :STATus1:FLAGs?, :STATus2:FLAGs? and :STATus3:FLAGs? only

4 :STATus3:FLAGs? only

## :MESSage Commands

These commands are used to exchange messages with a remote device (e.g. an optical amplifier) using the message forwarding facility. This facility allows commands and responses to be passed between the host computer and a remote device attached to the controller’s auxiliary serial port (Serial Port 1). Note that the facility is only available on FLM-0107, FLM-0130, FLM-0151 and FLM-0177 boards.

For further information on setting up the serial port and message forwarding, refer to Sections 5.1.3.4 and 5.1.3.2, respectively.

###### :MESSage <s>

This command is used to send a message (i.e. a command) to the remote device. The message to be sent should be specified as a string parameter but should exclude any terminator required by the remote device, so as not to conflict with the controller’s own command terminator. The controller can be programmed to append a terminator to the message when it sends it out to the remote device. Note that any lower-case characters contained in the message will be converted to upper-case prior to transmission to the remote device.

Syntax:	:MESSage <s>

Parameters:	<s>	message to remote device

Response:	none

###### :MESSage?

This command is used to read a message (i.e. a response) from the remote device. Characters received from the remote device are held in a buffer in the fiber laser controller until requested by this command. A maximum of 256 characters can be stored in the buffer.

Syntax:	:MESSage?

Parameters:	none

Response:	returns message from remote device

## :DIAGnostic Commands

These commands are used to control the diagnostic functions.

### :AMPLifier Commands

This group of commands is used to control the amplifier diagnostics. They are only available on the high power fiber laser systems (hardware application = “HPS-K” or “HPS-I”).

###### :ECHO <b>

This command is used to specify the echo state. When enabled, all commands and responses exchanged with the amplifier via the auxiliary (internal) serial port are also copied to the main serial port. This allows the amplifier communications to be monitored externally. Note that the controller needs to be unlocked in order to use this command.

Syntax:	:DIAGnostic:AMPLifier:ECHO <b>

Parameters:	<b>	echo state	0 = disabled; 1 = enabled

Configured:	no

Response:	none

###### :ECHO?

This command is used to query the echo state.

Syntax:	:DIAGnostic:AMPLifier:ECHO?

Parameters:	none

Response:	returns the echo state	0 = disabled; 1 = enabled

### :TEST Commands

This group of commands is used to control the PCB test functions.

###### :EEPRom?

This command is used to initiate the EEPROM test and query the result. It is only available on the FLM-0141 (Rev. D onwards), FLM-0151 and FLM-0177 boards. Note that the controller needs to be unlocked in order to use this command.

Syntax:	:DIAGnostic:TEST:EEPRom?

Parameters:	none

Response:	returns the EEPROM test result

The EEPROM comprises 32 pages, each of 16 bytes. The test writes and verifies 2 bit patterns at each address. The result is the number of byte/bit patterns that are verified successfully and should be 32 x 16 x 2 = 1024 if all locations are functioning correctly.

###### :LEDS <b>

This command is used to specify the LED test state. When enabled, the Temperature Stable and Second Mode LED’s will flash. Note that the controller needs to be unlocked in order to use this command.

Syntax:	:DIAGnostic:TEST:LEDS <b>

Parameters:	<b>	LED test state	0 = disabled; 1 = enabled

Configured:	no

Response:	none

###### :LEDS?

This command is used to query the LED test state.

Syntax:	:DIAGnostic:TEST:LEDS?

Parameters:	none

Response:	returns the LED test state	0 = disabled; 1 = enabled

###### :PADJust <b>

This command is used to specify the power adjust test state. It allows access to the power adjust commands, which normally can only be used when a board is configured as a high power source. Note that the controller needs to be unlocked in order to use this command.

Syntax:	:DIAGnostic:TEST:PADJust <b>

Parameters:	<b>	power adjust test state	0 = disabled; 1 = enabled

Configured:	no

Response:	none

###### :PADJust?

This command is used to query the power adjust test state.

Syntax:	:DIAGnostic:TEST:PADJust?

Parameters:	none

Response:	returns the power adjust test state	0 = disabled; 1 = enabled

###### :VREFerence <b>

This command is used to specify the reference voltage test state. When this feature is enabled, the 1.042 V and 2.048 V reference signals are smoothed using the 100-sample rolling averaging normally associated with the power and wavelength adjust inputs. The average values are read back using the :SYSTem:VREFerence<1|2>? command (see Section 5.1). These facilities are available on the FLM-0141 board only. Note that the controller needs to be unlocked in order to use this command.

Syntax:	:DIAGnostic:TEST:VREFerence <b>

Parameters:	<b>	Vref test state	0 = disabled; 1 = enabled

Configured:	no

Response:	none

###### :VREFerence?

This command is used to query the reference voltage test state. It is available on the FLM-0141 board only.

Syntax:	:DIAGnostic:TEST:VREFerence?

Parameters:	none

Response:	returns the Vref test state	0 = disabled; 1 = enabled

### :WDOG Commands

This group of commands is used to control the watchdog functions. For further details on using the functions, refer to the Watchdog Diagnositcs Commands document j.

###### :BOOT

This command uses the watchdog to reset the system and boot load via Serial Port 0. It is available on the FLM-0141, FLM-0151 and FLM-0177 boards only.

Syntax:	:DIAGnostic:WDOG:BOOT

Parameters:	none

Response:	none

###### :FLAG <b>

This command is used to clear the watchdog reset status flag.

Syntax:	:DIAGnostic:WDOG:FLAG <b>

Parameters:	<b>	action	0 = no effect; 1 = clears flag

Configured:	no

Response:	none

###### :FLAG?

This command is used to query the watchdog reset status flag. This flag, if set, indicates that the reset condition was generated by the watchdog. It can therefore be used to distinguish between the normal power-up reset and a watchdog reset. The flag remains latched until the user writes a 1 to clear the condition.

Syntax:	:DIAGnostic:WDOG:FLAG?

Parameters:	none

Response:	returns the reset status flag	0 = external (power-up) reset; 1 = watchdog reset

###### :INTerrupt <b>

This command is used to specify the watchdog interrupt enable state. If the interrupt is enabled, then the watchdog will generate an interrupt instead of resetting the system. The watchdog interrupt sends out a status report via the serial port – see Section 5.4.3.1 for further details. Note that the watchdog interrupt is always disabled at reset.

Syntax:	:DIAGnostic:WDOG:INTerrupt <b>

Parameters:	<b>	interrupt enable state	0 = disabled; 1 = enabled

Configured:	no

Response:	none

###### :INTerrupt?

This command is used to query the watchdog interrupt enable state.

Syntax:	:DIAGnostic:WDOG:INTerrupt?

Parameters:	none

Response:	returns the interrupt enable state	0 = disabled; 1 = enabled

###### :STATe <b>

This command is used to specify the watchdog enable state. Note that the watchdog is always disabled at reset and needs to be enabled before it can reset the system or generate in interrupt.

Syntax:	:DIAGnostic:WDOG:STATe <b>

Parameters:	<b>	enable state	0 = disabled; 1 = enabled

Configured:	no

Response:	none

###### :STATe?

This command is used to query the watchdog enable state.

Syntax:	:DIAGnostic:WDOG:STATe?

Parameters:	none

Response:	returns the enable state	0 = disabled; 1 = enabled

###### :TEST

This command is used to test the watchdog. It places the system into an infinite loop, causing the watchdog to reset the system or generate an interrupt, as appropriate.

Syntax:	:DIAGnostic:WDOG:TEST

Parameters:	none

Response:	none

#### Watchdog Status Report

The watchdog interrupt handler collects certain information to help identify the conditions immediately preceding the interrupt event. This includes the values of the stack pointer (SP) and return program counter (RPC) registers, the return address (from the interrupt handler) and various other CPU registers. The information is formatted into a 10-line status report, which is sent out on the serial port. An example of the status report is shown in Figure 5.3.

Figure 5.3 – Watchdog Status Report

WDINT Status Report

SP     : 0x0000008C

RPC    : 0x003EC6B1

T,  ST0: 0x00035485

AH, AL : 0x003E0001

PH, PL : 0x69D33B20

AR1,AR0: 0x83FC000E

DB, ST1: 0x01C02A08

DBG,IER: 0x001F0101

RET ADD: 0x003ECB05

## :OPTimization Commands

These commands are used to control the optimization functions. Note that they are only available on the FLM-0130, FLM-0141, FLM-0151 and FLM-0177 boards.

###### :DATA?

This command is used to query the raw power data acquired during the upward and downward ramp sequence. The data is returned using the IEEE 488.2 binary block format, which comprises a block header, followed by the data and, finally, the standard message terminator. The block header describes the size of the data block and has the following form:

#<number of digits><number of bytes>

The <number of digits> refers to the length of the <number of bytes> parameter, while the latter refers to the length of the data block itself. Each data point is a 16-bit unsigned integer and is transferred as two bytes, MSB first. The total number of data points will depend on the ramp settings. For example, with an initial dwell of 50, a step dwell of 5 and 30 steps in the ramp, the number of points will be 50  5  (30  2  1)  355.  This will be transmitted as:

#3710<MSB0><LSB0><MSB1><LSB1>...<MSB354><LSB354><terminator>

The data block is the array of power readings, in bits, recorded during the ramp sequence.

Syntax:	:OPTimization:DATA?

Parameters:	none

Response:	returns the raw power data

###### :FLAG?

This command is used to query the confidence flag. The flag indicates the reliability of the predicted optimum voltage. It is used in conjunction with the :SYSTem:OPTimization:CONFidence <i> setting (see Section 5.1.8) to assess whether or not to apply the predicted voltage.

Syntax:	:OPTimization:FLAG?

Parameters:	none

Response:	returns the confidence flag	see Table 5.32

Table 5.32 – Optimization Flag Values

| Value | Description |  |
| --- | --- | --- |
|  | f/w versions 5.4.003 to 6.2.001 | f/w version 6.3.000 onwards |
| 3 | Good prediction, based on quadratic fits to the ramp data | Good prediction, based on quadratic fits to the ramp data |
| 2 | Fairly good prediction, based on point of maximum power in the 3-point test | Fairly good prediction, based on widest single mode region |
| 1 | Questionable prediction, based on point of lowest second mode level in the 3-point test | Questionable prediction, based on widest single mode region being < 5 V wide |
| 0 | Very questionable prediction, based on all points being in 2nd mode in the 3-point test | Very questionable prediction, due to more than 2 second mode regions being found on either or both ramps |

###### :IDWell <i>

This command is used to specify the initial dwell time at the start of the ramp.

Syntax:	:OPTimization:IDWell <i>	(f/w versions 5.4.000 to 5.4.003 only)

Parameters:	<i>	initial dwell time in s/10

Configured:	yes

Response:	none

###### :IDWell?

This command is used to query the initial dwell time at the start of the ramp.

Syntax:	:OPTimization:IDWell? 		(f/w versions 5.4.000 to 5.4.003 only)

Parameters:	none

Response:	returns the initial dwell time in s/10

###### :POINt<0|1|2>?

This command is used to query the data from the 3-point test. The point number is specified as a suffix to the command header. Note that there should be no space between the header and the point number. The response message comprises two comma separated values. The first value represents the power readings accumulated at the point and the second represents the second mode level.

Syntax:	:OPTimization:POINT<0|1|2>?

Parameters:	none

Response:	returns the power value and second mode level in bits

###### :RMAX <f>

This command is used to specify the maximum ramp voltage.

Syntax:	:OPTimization:RMAX <f>		(f/w versions 5.4.001 to 5.4.003 only)

Parameters:	<f>	maximum voltage in V

Configured:	yes

Response:	none

###### :RMAX?

This command is used to query the maximum ramp voltage.

Syntax:	:OPTimization:RMAX?		(f/w versions 5.4.001 to 5.4.003 only)

Parameters:	none

Response:	returns the maximum voltage in V

###### :RMIN <f>

This command is used to specify the minimum ramp voltage.

Syntax:	:OPTimization:RMIN <f>		(f/w versions 5.4.001 to 5.4.003 only)

Parameters:	<f>	minimum voltage in V

Configured:	yes

Response:	none

###### :RMIN?

This command is used to query the minimum ramp voltage.

Syntax:	:OPTimization:RMIN? 		(f/w versions 5.4.001 to 5.4.003 only)

Parameters:	none

Response:	returns the minimum voltage in V

###### :RSTEPs <i>

This command is used to specify the number of steps on the ramp.

Syntax:	:OPTimization:RSTeps <i>	(f/w versions 5.4.000 to 5.4.003 only)

Parameters:	<i>	number of steps

Configured:	yes

Response:	none

###### :RSTEPs?

This command is used to query the number of steps on the ramp.

Syntax:	:OPTimization:RSTeps?		(f/w versions 5.4.000 to 5.4.003 only)

Parameters:	none

Response:	returns the number of steps

###### :SDWell <i>

This command is used to specify the dwell time at each step on the ramp.

Syntax:	:OPTimization:SDWell <i>	(f/w versions 5.4.000 to 5.4.003 only)

Parameters:	<i>	step dwell time in s/10

Configured:	yes

Response:	none

###### :SDWell?

This command is used to query the dwell time at each step on the ramp.

Syntax:	:OPTimization:SDWell?		(f/w versions 5.4.000 to 5.4.003 only)

Parameters:	none

Response:	returns the step dwell time in s/10

###### :SMAX <i>

This command is used to specify the maximum step position.

Syntax:	:OPTimization:SMAX <i>		(f/w version 5.4.000 only)

Parameters:	<i>	maximum step position in V

Configured:	yes

Response:	none

###### :SMAX?

This command is used to query the maximum step position.

Syntax:	:OPTimization:SMAX?		(f/w version 5.4.000 only)

Parameters:	none

Response:	returns the maximum step position in V

###### :SMIN <i>

This command is used to specify the minimum step position.

Syntax:	:OPTimization:SMIN <i>		(f/w version 5.4.000 only)

Parameters:	<i>	minimum step position in V

Configured:	yes

Response:	none

###### :SMIN?

This command is used to query the minimum step position.

Syntax:	:OPTimization:SMIN?		(f/w version 5.4.000 only)

Parameters:	none

Response:	returns the minimum step position in V

###### :STATe <b>

This command is used to specify the optimization state.

Syntax:	:OPTimization:STATe <b>

Parameters:	<b>	optimization state	0 = stop optimization; 1 = start optimization

Configured:	yes

Response:	none

###### :STATe?

This command is used to query the optimization state.

Syntax:	:OPTimization:STATe?

Parameters:	none

Response:	returns the optimization state	0 = optimization stopped; 1 = optimization running

###### :VOPTimum?

This command is used to query the optimum voltage.

Syntax:	:OPTimization:VOPTimum?

Parameters:	none

Response:	returns the optimum voltage

### :LOG Commands

These commands are used to access the optimization log. They are only available on the FLM-0141 (Rev. D onwards), FLM-0151 and FLM-0177 boards.

###### :CLEar

This command is used to clear the optimization log. It uses a special password protection mechanism and is reserved for factory use only.

Syntax:	:OPTimization:LOG:CLEar

Parameters:	none

Response:	none

###### :INDex?

This command is used to query the index of the last optimization event.

Syntax:	:OPTimization:LOG:INDex?

Parameters:	none

Response:	returns the index of the last event

###### :TOTal?

This command is used to query the total number of optimization events.

Syntax:	:OPTimization:LOG:TOTal?

Parameters:	none

Response:	returns the total number of events

#### :EVENt<i> Commands

This group of commands is used to query the optimization event data. Optimization events are logged into a circular buffer. Up to 31 events can be accommodated, each comprising the time, heater temperature, optimized piezo voltage and confidence flag. The event number, which should be in the range 1 to 31, is specified as a suffix to the command header. Note that there should be no space between the header and the suffix.

###### :FLAG?

This command is used to query the confidence flag for the specified optimization event.

Syntax:	:OPTimization:LOG:EVENt<i>:FLAG?

Parameters:	none

Response:	returns the confidence flag

###### :TEMPerature?

This command is used to query the heater temperature for the specified optimization event.

Syntax:	:OPTimization:LOG:EVENt<i>:TEMPerature?

Parameters:	none

Response:	returns the temperature in °C

###### :TIME?

This command is used to query the time of the specified optimization event.

Syntax:	:OPTimization:LOG:EVENt<i>:TIME?

Parameters:	none

Response:	returns the time in hours

###### :VOLTage?

This command is used to query the optimized piezo voltage for the specified optimization event.

Syntax:	:OPTimization:LOG:EVENt<i>:VOLTage?

Parameters:	none

Response:	returns the voltage in V

### :RAMP<0|1> Commands

This group of commands is used to query the ramp data results. The upward (0) and downward (1) ramp is specified as a suffix to the command header. Note that there should be no space between the header and the ramp number.

###### :DIPS?

This command is used to query the power dip positions. The dip location process accommodates up to two dips on each ramp. The response message comprises up to three comma separated values. The first value represents the number of dips identified and will normally be 0, 1 or 2. The remaining values, if sent, are the dip positions and refer to the step number on the ramp.

Syntax:	:OPTimization:RAMP<0|1>:DIPS?

Parameters:	none

Response:	returns the dip positions

###### :PEAKs?

This command is used to query the power peak positions. The peak location process accommodates up to two peaks on each ramp. The response message comprises up to three comma separated values. The first value represents the number of peaks identified and will be 0, 1 or 2. The remaining values, if sent, are the peak positions in V.

Syntax:	:OPTimization:RAMP<0|1>:PEAKs?

Parameters:	none

Response:	returns the peak positions

###### :POWers?

This command is used to query the power readings accumulated at each step on the ramp. The response message is a series of comma separated values. Note that the number of values is always one greater than the number of steps.

Syntax:	:OPTimization:RAMP<0|1>:POWers?	(f/w version 5.4.007 onwards)

Parameters:	none

Response:	returns the power values in bits

###### :QUADratic<0|1>?

This command is used to query the quadratic fit results. Up to two quadratic fits are calculated for each ramp, based on the data segments on each side of the power dips. The response message comprises up to five comma separated values, depending on firmware version. The first value (on version 5.4.004 onwards) represents the number of points in the data segment, the next three values are the 0th, 1st and 2nd order quadratic coefficients and the final value (on version 5.4.006 onwards) is the coefficient of determination. The latter provides a measure of the fit quality and lies in the range 0 (poor) to 1 (good). It is used in conjunction with the :SYSTem:OPTimization:QUADratic:RMIN <f> setting (see Section 5.1.8.3) to validate the fit.

Syntax:	:OPTimization:RAMP<0|1>:QUADratic<0|1>?

Parameters:	none

Response:	returns the quadratic fit data

###### :SMLevels?

This command is used to query the second mode levels accumulated at each step on the ramp. The response message is a series of comma separated values. Note that the number of values is always one greater than the number of steps.

Syntax:	:OPTimization:RAMP<0|1>:SMLevels?	(f/w version 5.4.007 onwards)

Parameters:	none

Response:	returns the second mode levels in bits

###### :SMRanges?

This command is used to query the positions of the second mode ranges. The second mode location process accommodates up to two ranges on each ramp. The response message comprises up to five comma separated values. The first value represents the number of ranges identified and will normally be 0, 1 or 2. The remaining values, if sent, represent the beginning and end points of each range and refer to the step number on the ramp.

Syntax:	:OPTimization:RAMP<0|1>:SMRanges?

Parameters:	none

Response:	returns the second mode ranges

###### :STEPs?

This command is used to query the power readings accumulated at each step on the ramp. The response message is a series of comma separated values. Note that the number of values is always one greater than the number of steps.

Syntax:	:OPTimization:RAMP<0|1>:STEPs?	(f/w versions 5.4.003 to 5.4.006 only)

Parameters:	none

Response:	returns the step data values in bits

Table 5.33 – SYSTem Command Summary

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :SYSTem | Commands to set up/query system settings: |  |  | 1.0.000 |
| :ADCOffset <i> | Specify ADC offset | bits |  | 1.0.003 |
| :ADCOffset? | Query ADC offset | bits |  | 1.0.003 |
| :AMPLifier | Path to set up/query amplifier settings: |  |  | 7.5.000 |
| :C[0¦1¦2]0 <f> | Specify current coefficient C[0¦1¦2]0 | see text |  | see text |
| :C[0¦1¦2]0? | Query current coefficient C[0¦1¦2]0 | see text |  | see text |
| :C[0¦1¦2]1 <f> | Specify current coefficient C[0¦1¦2]1 | see text |  | see text |
| :C[0¦1¦2]1? | Query current coefficient C[0¦1¦2]1 | see text |  | see text |
| :CMAXimum<[0]¦1> <f> | Specify maximum current set-point | A |  | see text |
| :CMAXimum<[0]¦1>? | Query maximum current set-point | A |  | see text |
| :CMINimum<[0]¦1> <f> | Specify minimum current set-point | A |  | see text |
| :CMINimum<[0]¦1>? | Query minimum current set-point | A |  | see text |
| :CMONitor? | Query current monitor | A |  | 8.1.000 |
| :CONTrol | Path to set up/query power control settings: |  |  | 7.5.000 |
| :KD <f> | Specify PID differential constant | bits/W.s |  | 7.5.000 |
| :KD? | Query PID differential constant | bits/W.s |  | 7.5.000 |
| :KI <f> | Specify PID integral constant | bits/W/s |  | 7.5.000 |
| :KI? | Query PID integral constant | bits/W/s |  | 7.5.000 |
| :KP <f> | Specify PID proportional constant | bits/W |  | 7.5.000 |
| :KP? | Query PID proportional constant | bits/W |  | 7.5.000 |
| :CSETpoint<[0]¦1> <f> | Specify current set-point | A |  | see text |
| :CSETpoint<[0]¦1>? | Query current set-point | A |  | see text |
| :FLAGs? | Query status flags |  |  | 7.5.000 |
| :MSELect <i> | Specify operating mode (0 = APC, 1 = ACC) |  |  | 7.5.000 |
| :MSELect? | Query operating mode |  |  | 7.5.000 |
| :OPTions <i> | Specify options |  |  | 8.4.002 |
| :OPTions? | Query options |  |  | 8.4.002 |
| :P[0¦1]0 <f> | Specify power monitor coefficient P[0¦1]0 | mW |  | see text |
| :P[0¦1]0? | Query power monitor coefficient P[0¦1]0 | mW |  | see text |
| :P[0¦1]1 <f> | Specify power monitor coefficient P[0¦1]1 | mW/bit |  | see text |
| :P[0¦1]1? | Query power monitor coefficient P[0¦1]1 | mW/bit |  | see text |
| :PMAXimum <f> | Specify maximum power set-point | W |  | 7.5.000 |
| :PMAXimum? | Query maximum power set-point | W |  | 7.5.000 |
| :PMINimum <f> | Specify minimum power set-point | W |  | 7.5.000 |
| :PMINimum? | Query minimum power set-point | W |  | 7.5.000 |
| :PMONitor<[0]¦1>? | Query power monitor | see text |  | 7.5.000 |
| :PSETpoint <f> | Specify power set-point | W |  | 7.5.000 |
| :PSETpoint? | Query power set-point | W |  | 7.5.000 |
| :RRATe<[0]¦1> <f> | Specify current set-point ramp rate | A/s/10 |  | see text |
| :RRATe<[0]¦1>? | Query current set-point ramp rate | A/s/10 |  | see text |
| :STATe <i> | Specify output state (0 = disabled, 1 = enabled) |  |  | 7.5.000 |
| :STATe? | Query output state |  |  | 7.5.000 |
| :TFACtor[0¦1] <f> | Specify power monitor tap factor |  |  | see text |
| :TFACtor[0¦1]? | Query power monitor tap factor |  |  | see text |
| :THReshold<0¦1> <f> | Specify power threshold | mW |  | 7.5.000 |
| :THReshold<0¦1>? | Query power threshold | mW |  | 7.5.000 |
| :TMONitor<0¦1>? | Query temperature monitor | °C |  | 8.1.000 |
| :BFACet[0¦1] | Path to set up/query back facet settings: |  |  | see text |
| :FLAGs? | Query status flags |  |  | 3.0.000 |
| :LEVEl? | Query back facet level | bits |  | 3.0.000 |
| :THReshold <i> | Specify back facet threshold | bits |  | 3.0.000 |
| :THReshold? | Query back facet threshold | bits |  | 3.0.000 |
| :COMMunicate | Path to set up/query communication settings: |  |  | 4.4.000 |
| :ADDRess | Path to set up/query address settings: |  |  | 6.0.003 |
| :ALIas <s> | Specify alias |  | ü | 6.0.003 |
| :ALIas? | Query alias |  |  | 6.0.003 |
| :PHYSical? | Query physical (hardware) address |  |  | 6.0.003 |
| :STATe <b> | Specify addressing state (0 = disabled, 1 = enabled) |  | ü | 6.0.003 |
| :STATe? | Query addressing state |  |  | 6.0.003 |
| :FORWard | Path to set up/query message forwarding settings: |  |  | 4.4.000 |
| :STATe <b> | Specify forwarding state (0 = disabled, 1 = enabled) |  |  | 4.4.000 |
| :STATe? | Query forwarding state |  |  | 4.4.000 |
| :TERMinator<0¦1> <i> | Specify forwarding terminator |  |  | 4.4.000 |
| :TERMinator<0¦1>? | Query forwarding terminator |  |  | 4.4.000 |
| :SERial<0> | Path to set up/query Serial Port 0 settings: |  |  | 5.0.000 |
| :MODE <i> | Specify mode (0 = normal, 1 = listen only) |  |  | 5.0.000 |
| :MODE? | Query mode |  |  | 5.0.000 |
| :SERial<1> | Path to set up/query Serial Port 1 settings: |  |  | 4.4.000 |
| :BAUD <i> | Specify baud rate |  |  | 4.4.000 |
| :BAUD? | Query baud rate |  |  | 4.4.000 |
| :BITS <i> | Specify number of data bits |  |  | 4.4.000 |
| :BITS? | Query number of data bits |  |  | 4.4.000 |
| :PARity <i> | Specify parity |  |  | 4.4.000 |
| :PARity? | Query parity |  |  | 4.4.000 |
| :SBITs <i> | Specify number of stop bits |  |  | 4.4.000 |
| :SBITs? | Query number of stop bits |  |  | 4.4.000 |

Table 5.33 – SYSTem Command Summary (continued)

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :ERRor? | Query (read) next error and remove from queue |  |  | 1.0.000 |
| :FIRMware | Path to query firmware details: |  |  | 1.0.000 |
| :DATE? | Query firmware release date |  |  | 1.0.000 |
| :VERSion? | Query firmware version |  |  | 1.0.000 |
| :HARDware | Path to set up/query hardware settings: |  |  | 1.0.000 |
| :APPLication <s> | Specify application |  |  | 4.2.000 |
| :APPLication? | Query application |  |  | 4.2.000 |
| :DATE <y>,<m>,<d> | Specify manufacture date |  |  | 1.0.000 |
| :DATE? | Query manufacture date |  |  | 1.0.000 |
| :LCDWelcome <s> | Specify LCD welcome message |  |  | 4.5.001 |
| :LCDWelcome? | Query LCD welcome message |  |  | 4.5.001 |
| :MANufacturer <s> | Specify manufacturer |  |  | 8.1.000 |
| :MANufacturer? | Query manufacturer |  |  | 8.1.000 |
| :MODel <s> | Specify model number |  |  | 8.1.000 |
| :MODel? | Query model number |  |  | 8.1.000 |
| :OPTions <i> | Specify PCB options |  |  | 2.0.000 |
| :OPTions? | Query PCB options |  |  | 2.0.000 |
| :REVision <a> | Specify PCB revision |  |  | 2.0.000 |
| :REVision? | Query PCB revision |  |  | 2.0.000 |
| :SERial <s> | Specify serial number |  |  | 1.0.000 |
| :SERial? | Query serial number |  |  | 1.0.000 |
| :TYPE <i> | Specify PCB type |  |  | 2.0.000 |
| :TYPE? | Query PCB type |  |  | 2.0.000 |
| :MHFTuning | Path to set up/query mode hop free tuning settings: |  |  | 1.0.001 |
| :BAND<a¦i> | Path to set up/query mode hop free band settings: |  |  | see text |
| :HALFwidth <f> | Specify band halfwidth | pm |  | 5.0.000 |
| :HALFwidth? | Query band halfwidth | pm |  | 5.0.000 |
| :NBSLope <f> | Specify temperature sensitivity of narrow-band grating | pm/ºC |  | 5.0.000 |
| :NBSLope? | Query temperature sensitivity of narrow-band grating | pm/ºC |  | 5.0.000 |
| :NBTemp <f> | Specify temperature of narrow-band grating | ºC |  | 1.0.001 |
| :NBTemp? | Query temperature of narrow-band grating | ºC |  | 1.0.001 |
| :RTEMperature <f> | Specify reference temperature | ºC |  | 5.0.000 |
| :RTEMperature? | Query reference temperature | ºC |  | 5.0.000 |
| :RVOLtage <f> | Specify reference voltage | V |  | 5.0.000 |
| :RVOLtage? | Query reference voltage | V |  | 5.0.000 |
| :RWAVelength <f> | Specify reference wavelength | nm |  | 1.0.001 |
| :RWAVelength? | Query reference wavelength | nm |  | 1.0.001 |
| :TSLope <f> | Specify temperature sensitivity | pm/ºC |  | 5.0.000 |
| :TSLope? | Query temperature sensitivity | pm/ºC |  | 5.0.000 |
| :VSLope <f> | Specify voltage sensitivity | pm/V |  | 5.0.000 |
| :VSLope? | Query voltage sensitivity | pm/V |  | 5.0.000 |
| :WBHalfrange <f> | Specify half-range of wide-band grating | ºC |  | see text |
| :WBHalfrange? | Query half-range of wide-band grating | ºC |  | see text |
| :WBSLope <f> | Specify temperature sensitivity of wide-band grating | pm/ºC |  | 5.0.000 |
| :WBSLope? | Query temperature sensitivity of wide-band grating | pm/ºC |  | 5.0.000 |
| :WBTemp <f> | Specify temperature of wide-band grating | ºC |  | 1.0.001 |
| :WBTemp? | Query temperature of wide-band grating | ºC |  | 1.0.001 |
| :NBSLope <f> | Specify temperature sensitivity of narrow-band grating | pm/ºC |  | see text |
| :NBSLope? | Query temperature sensitivity of narrow-band grating | pm/ºC |  | see text |
| :WBSLope <f> | Specify temperature sensitivity of wide-band grating | pm/ºC |  | see text |
| :WBSLope? | Query temperature sensitivity of wide-band grating | pm/ºC |  | see text |
| :OPTimzation | Path to set up/query optimization settings: |  |  | 5.4.003 |
| :CONFidence <i> | Specify minimum confidence level |  |  | 5.4.004 |
| :CONFidence? | Query minimum confidence level |  |  | 5.4.004 |
| :DIP | Path to set up/query dip detection settings: |  |  | 5.4.003 |
| :DMIN <f> | Specify minimum depth |  |  | 5.4.003 |
| :DMIN? | Query minimum depth |  |  | 5.4.003 |
| :SMIN <f> | Specify minimum spacing | V |  | 5.4.003 |
| :SMIN? | Query minimum spacing | V |  | 5.4.003 |
| :STATe <b> | Specify detection state (0 = disabled, 1 = enabled) |  |  | 5.4.003 |
| :STATe? | Query detection state |  |  | 5.4.003 |
| :WMAX <f> | Specify maximum width | V |  | 5.4.003 |
| :WMAX? | Query maximum width | V |  | 5.4.003 |
| :PEAK | Path to set up/query peak settings: |  |  | 5.4.003 |
| :PMAX<0¦1> <f> | Specify maximum peak voltage | V |  | 5.4.003 |
| :PMAX<0¦1>? | Query maximum peak voltage | V |  | 5.4.003 |
| :PMIN<0¦1> <f> | Specify minimum peak voltage | V |  | 5.4.003 |
| :PMIN<0¦1>? | Query minimum peak voltage | V |  | 5.4.003 |
| :SMAX <f> | Specify maximum peak separation | V |  | 5.4.003 |
| :SMAX? | Query maximum peak separation | V |  | 5.4.003 |
| :SMIN <f> | Specify minimum peak separation | V |  | 5.4.003 |
| :SMIN? | Query minimum peak separation | V |  | 5.4.003 |

Table 5.33 – SYSTem Command Summary (continued)

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :QUADratic | Path to set up/query quadratic fit settings: |  |  | 5.4.004 |
| :CMAX <f> | Specify maximum 2nd order coefficient |  |  | 5.4.004 |
| :CMAX? | Query maximum 2nd order coefficient |  |  | 5.4.004 |
| :CMIN <f> | Specify minimum 2nd order coefficient |  |  | 5.4.004 |
| :CMIN? | Query minimum 2nd order coefficient |  |  | 5.4.004 |
| :NMIN <i> | Specify minimum number of points |  |  | 5.4.004 |
| :NMIN? | Query minimum number of points |  |  | 5.4.004 |
| :RMIN <f> | Specify minimum coefficient of determination |  |  | 5.4.006 |
| :RMIN? | Query minimum coefficient of determination |  |  | 5.4.006 |
| :RAMP | Path to set up/query ramp settings: |  |  | 5.4.003 |
| :IDWell <i> | Specify initial dwell time | s/10 |  | 5.4.003 |
| :IDWell? | Query initial dwell time | s/10 |  | 5.4.003 |
| :SDWell <i> | Specify step dwell time | s/10 |  | 5.4.003 |
| :SDWell? | Query step dwell time | s/10 |  | 5.4.003 |
| :STEPs <i> | Specify number of steps |  |  | 5.4.003 |
| :STEPs? | Query number of steps |  |  | 5.4.003 |
| :VMAX <f> | Specify maximum voltage | V |  | 5.4.003 |
| :VMAX? | Query maximum voltage | V |  | 5.4.003 |
| :VMIN <f> | Specify minimum voltage | V |  | 5.4.003 |
| :VMIN? | Query minimum voltage | V |  | 5.4.003 |
| :SAVE <b> | Specify auto-save state |  |  | 5.4.004 |
| :SAVE? | Query auto-save state |  |  | 5.4.004 |
| :PADJust | Path to set up/query power adjust settings: |  |  | 4.8.000 |
| :INPut <i> | Specify input (0 = BNC socket, 1 = potentiometer) |  |  | 4.8.000 |
| :INPut? | Query input |  |  | 4.8.000 |
| :PMAX <f> | Specify maximum power | mW |  | 4.8.000 |
| :PMAX? | Query maximum power | mW |  | 4.8.000 |
| :PMIN <f> | Specify minimum power | mW |  | 4.8.000 |
| :PMIN? | Query minimum power | mW |  | 4.8.000 |
| :POWer <f> | Specify power | mW |  | 4.8.000 |
| :POWer? | Query power | mW |  | 4.8.000 |
| :MUTE <b> | Specify mute state (0 = disabled, 1 = enabled) |  |  | 7.0.000 |
| :MUTE? | Query mute state |  |  | 7.0.000 |
| :STATe <i> | Specify adjustment state (0 = disabled, 1 = enabled) |  |  | 4.8.000 |
| :STATe? | Query adjustment state |  |  | 4.8.000 |
| :PASSword | Path to set up/query password settings: |  |  | 1.0.000 |
| :CDISable <s> | Disable protected commands |  |  | 1.0.000 |
| :CENable <s> | Enable protected commands |  |  | 1.0.000 |
| :STATe? | Query enable state (0 = disabled, 1 = enabled) |  |  | 1.0.000 |
| :NEW <s1>,<s2> | Specify new password |  |  | 1.0.000 |
| :PCONtrol | Path to set up/query pump control settings: |  |  | 1.0.001 |
| :CLIMit[0¦1] <i> | Specify current limit | mA |  | see text |
| :CLIMit[0¦1]? | Query current limit | mA |  | see text |
| :DELay <b> | Specify startup delay state (0 = disabled, 1 = enabled) |  |  | 5.2.000 |
| :DELay? | Query startup delay state |  |  | 5.2.000 |
| :EMODulation <b> | Specify ext. modulation state (0 = disabled, 1 = enabled) |  |  | 7.0.000 |
| :EMODulation? | Query ext. modulation state |  |  | 7.0.000 |
| :FLAGs? | Query status flags |  |  | 1.0.001 |
| :MODE<0¦1> | Path to set up/query operating mode settings: |  |  | 1.0.001 |
| :KD <f> | Specify PID differential constant Mode 0 only | bits/mW.s |  | 1.1.000 |
| :KD? | Query PID differential constant Mode 0 only | bits/mW.s |  | 1.1.000 |
| :KI <f> | Specify PID integral constant Mode 0 only | bits/mW/s |  | 1.1.000 |
| :KI? | Query PID integral constant Mode 0 only | bits/mW/s |  | 1.1.000 |
| :KP <f> | Specify PID proportional constant Mode 0 only | bits/mW |  | 1.1.000 |
| :KP? | Query PID proportional constant Mode 0 only | bits/mW |  | 1.1.000 |
| :RANGe <i> | Specify range |  |  | 1.0.001 |
| :RANGe? | Query range |  |  | 1.0.001 |
| :SPOint <i> | Specify set-point | see text |  | 1.0.001 |
| :SPOint[0¦1] <f> | Specify set-point | see text |  | see text |
| :SPOint[0¦1]? | Query set-point | see text |  | see text |
| :MSELect <i> | Specify mode select method |  |  | 1.0.001 |
| :MSELect? | Query mode select method |  |  | 1.0.001 |
| :PLIMit <i> | Specify power limit | mW |  | Note 1 |
| :PLIMit <f> | Specify power limit | mW |  | 6.2.000 |
| :PLIMit? | Query power limit | mW |  | Note 2 |
| :STATe <b> | Specify power limiting state (0 = disabled, 1 = enabled) |  |  | 6.2.000 |
| :STATe? | Query power limiting state |  |  | 6.2.000 |
| :PMODe <b> | Specify power mode state (0 = disabled, 1 = enabled) |  |  | 7.4.000 |
| :PMODe? | Query power mode state |  |  | 7.4.000 |
| :RANGe<0[¦1¦2¦3]> | Path to set up/query range settings: |  |  | 1.0.001 |
| :CMAX[0¦1] <i> | Specify maximum current | mA |  | see text |
| :CMAX[0¦1]? | Query maximum current | mA |  | see text |
| :RRATe<0¦1> <i> | Specify ramp rate | mA/s/10 |  | 8.5.000 |
| :RRATe<0¦1>? | Query ramp rate | mA/s/10 |  | 8.5.000 |

Notes

1. Supported by version 5.4.012 to 5.4.014 only.

2. Supported by version 5.4.012 to 5.4.014 and version 6.2.000 onwards.

Table 5.33 – SYSTem Command Summary (continued)

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :SPOint <i> | Specify set-point | see text |  | see text |
| :SPOint <f> | Specify set-point | see text |  | see text |
| :SPOint? | Query set-point | see text |  | see text |
| :STATe <b> | Specify pump state (0 = disabled, 1 = enabled) |  |  | 5.0.002 |
| :STATe? | Query pump state |  |  | 5.0.002 |
| :PDRive | Path to set up/query piezo drive settings: |  |  | 5.0.000 |
| :MODulation | Path to set up/query modulation settings: |  |  | 5.0.000 |
| :AMPLitude <f> | Specify amplitude | V |  | 5.0.000 |
| :AMPLitude? | Query amplitude | V |  | 5.0.000 |
| :FREQuency <f> | Specify frequency | Hz |  | 5.0.000 |
| :FREQuency? | Query frequency | Hz |  | 5.0.000 |
| :STATe <b> | Specify modulation state (0 = disabled, 1 = enabled) |  |  | 5.0.000 |
| :STATe? | Query modulation state |  |  | 5.0.000 |
| :SFACtor <f> | Specify scale factor |  |  | 6.5.000 |
| :SFACtor? | Query scale factor |  |  | 6.5.000 |
| :VOLTage <f> | Specify DC voltage | V |  | 5.0.000 |
| :VOLTage? | Query DC voltage | V |  | 5.0.000 |
| :PMONitor | Path to set up/query pump monitor settings: |  |  | 1.0.001 |
| :C[0¦1]0 <f> | Specify current coefficient C[0¦1]0 | mA |  | see text |
| :C[0¦1]0? | Query current coefficient C[0¦1]0 | mA |  | see text |
| :C[0¦1]1 <f> | Specify current coefficient C[0¦1]1 | mA/bit |  | see text |
| :C[0¦1]1? | Query current coefficient C[0¦1]1 | mA/bit |  | see text |
| :CURRent[0¦1]? | Query monitored current | mA |  | see text |
| :P[0¦1]0 <f> | Specify power coefficient P[0¦1]0 | mW |  | see text |
| :P[0¦1]0? | Query power coefficient P[0¦1]0 | mW |  | see text |
| :P[0¦1]1 <f> | Specify power coefficient P[0¦1]1 | mW/bit |  | see text |
| :P[0¦1]1? | Query power coefficient P[0¦1]1 | mW/bit |  | see text |
| :POWer[0¦1]? | Query monitored power | mW |  | see text |
| :TFACtor[0¦1] <f> | Specify output tap factor |  |  | see text |
| :TFACtor[0¦1]? | Query output tap factor |  |  | see text |
| :RIN | Path to set up/query RIN suppression settings: |  |  | 1.0.000 |
| :ADJust | Path to set up/query automatic adjustment settings: |  |  | 6.4.001 |
| :FLAGs? | Query status flags |  |  | 6.4.001 |
| :GAIN? | Query active gain setting |  |  | 6.4.001 |
| :MONitor? | Query RIN monitor level |  |  | 6.4.001 |
| :STATe <b> | Specify auto-adjust state (0 = disabled, 1 = enabled) |  |  | 6.4.001 |
| :STATe? | Query auto-adjust state |  |  | 6.4.001 |
| :APASs <i> | Specify all pass network setting |  |  | 6.0.000 |
| :APASs? | Query all pass network setting |  |  | 6.0.000 |
| :FLAGs? | Query status flags |  |  | 1.0.001 |
| :GAIN <i> | Specify gain setting |  |  | 5.0.000 |
| :GAIN? | Query gain setting |  |  | 5.0.000 |
| :HPASs <i> | Specify high pass filter setting |  |  | 6.0.000 |
| :HPASs? | Query high pass filter setting |  |  | 6.0.000 |
| :MSELect <i> | Specify mode select method |  |  | 1.0.000 |
| :MSELect? | Query mode select method |  |  | 1.0.000 |
| :RTIMe<0> 0 | Reset run time counter |  |  | 6.5.000 |
| :RTIMe<0¦1>? | Query run time counter |  |  | 6.5.000 |
| :SMODe | Path to set up/query second mode settings: |  |  | 1.0.000 |
| :CORRection | Path to set up/query second mode correction settings: |  |  | 4.2.000 |
| :FLAGs? | Query status flags |  |  | 5.0.000 |
| :HOLDoff <i> | Specify hold-off time | seconds |  | 5.2.001 |
| :HOLDoff? | Query hold-off time | seconds |  | 5.2.001 |
| :RAMP | Path to set up/query ramp settings: |  |  | 5.0.000 |
| :AMPLitude <f> | Specify amplitude | V |  | 5.0.000 |
| :AMPLitude? | Query amplitude | V |  | 5.0.000 |
| :DWELl <i> | Specify dwell time | s/10 |  | 5.0.000 |
| :DWELl? | Query dwell time | s/10 |  | 5.0.000 |
| :RANGe <i> | Specify range | steps |  | 4.6.000 |
| :RANGe? | Query range | steps |  | 4.6.000 |
| :STATe <b> | Specify correction state (0 = disabled, 1 = enabled) |  |  | Note 1 |
| :STATe <i> | Specify adjustment state |  |  | 6.3.000 |
| :STATe? | Query correction state |  |  | 4.2.000 |
| :STEP <f> | Specify step size | see text |  | 4.2.000 |
| :STEP? | Query step size | see text |  | 4.2.000 |
| :FLAGs? | Query status flags |  |  | 1.0.001 |
| :LEDDisable <b> | Specify LED disable state (1 = disabled, 0 = enabled) |  |  | 5.0.000 |
| :LEDDisable? | Query LED disable state |  |  | 5.0.000 |
| :LEVEl? | Query second mode level | bits |  | 1.0.000 |
| :THReshold <i> | Specify second mode threshold | bits |  | 1.0.000 |
| :THReshold? | Query second mode threshold | bits |  | 1.0.000 |

Notes

1. Supported by version 4.2.000 to 6.2.001 only.

Table 5.33 – SYSTem Command Summary (continued)

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :TCONtrol<0[¦1¦2¦3¦4]> | Path to set up/query temperature control settings: |  |  | see text |
| :DELay <i> | Specify heater startup delay | s |  | 6.5.001 |
| :DELay? | Query heater startup delay | s |  | 6.5.001 |
| :DOUTput <i> | Specify heater output value directly |  |  | 5.1.002 |
| :DOUTput? | Query heater output value |  |  | 5.1.002 |
| :ELEMents<0¦1> <i> | Specify heater element configuration |  |  | 5.1.000 |
| :ELEMents<0¦1>? | Query heater element configuration |  |  | 5.1.000 |
| :FLAGs? | Query status flags |  |  | 1.0.000 |
| :KD <f> | Specify PID differential constant | bits/ºC.s |  | 1.0.000 |
| :KD? | Query PID differential constant | bits/ºC.s |  | 1.0.000 |
| :KI <f> | Specify PID integral constant | bits/ºC/s |  | 1.0.000 |
| :KI? | Query PID integral constant | bits/ºC/s |  | 1.0.000 |
| :KP <f> | Specify PID proportional constant | bits/ºC |  | 1.0.000 |
| :KP? | Query PID proportional constant | bits/ºC |  | 1.0.000 |
| :OMAX <f> | Specify maximum output value | bits |  | 8.1.000 |
| :OMAX? | Query maximum output value | bits |  | 8.1.000 |
| :OMIN <f> | Specify minimum output value | bits |  | 8.1.000 |
| :OMIN? | Query minimum output value | bits |  | 8.1.000 |
| :PWMPeriod <i> | Specify heater PWM period | bits |  | 6.0.000 |
| :PWMPeriod? | Query heater PWM period | bits |  | 6.0.000 |
| :SPOint <f> | Specify set-point | ºC |  | 1.0.001 |
| :SPOint? | Query set-point | ºC |  | 1.0.001 |
| :STATe <b> | Specify control state (0 = disabled, 1 = enabled) |  |  | 1.0.000 |
| :STATe? | Query control state |  |  | 1.0.000 |
| :TMAX <f> | Specify over-temperature alarm limit | ºC |  | 1.0.000 |
| :TMAX? | Query over-temperature alarm limit | ºC |  | 1.0.000 |
| :TMIN <f> | Specify under-temperature alarm limit | ºC |  | 1.0.000 |
| :TMIN? | Query under-temperature alarm limit | ºC |  | 1.0.000 |
| :TOLerance <f> | Specify tolerance | ºC |  | 1.0.001 |
| :TOLerance? | Query tolerance | ºC |  | 1.0.001 |
| :VMONitor[1¦2¦3]? | Query heater voltage monitor | V |  | see text |
| :TMONitor<0[¦1¦2¦3¦4]> | Path to set up/query temperature monitor settings: |  |  | see text |
| :FLAGs? | Query status flags |  |  | 1.0.001 |
| :T0 <f> | Specify temperature coefficient T0 | ºC |  | 1.0.000 |
| :T0? | Query temperature coefficient T0 | ºC |  | 1.0.000 |
| :T1 <f> | Specify temperature coefficient T1 | ºC/bit |  | 1.0.000 |
| :T1? | Query temperature coefficient T1 | ºC/bit |  | 1.0.000 |
| :T2 <f> | Specify temperature coefficient T2 | ºC/bit² |  | 1.0.000 |
| :T2? | Query temperature coefficient T2 | ºC/bit² |  | 1.0.000 |
| :T3 <f> | Specify temperature coefficient T3 | ºC/bit³ |  | 1.0.000 |
| :T3? | Query temperature coefficient T3 | ºC/bit³ |  | 1.0.000 |
| :TEMPerature? | Query monitored temperature | ºC |  | 1.0.000 |
| :VOA | Path to set up/query variable optical attenuator settings: |  |  | 7.0.000 |
| :ATTenuation <f> | Specify attenuation set-point |  |  | 7.0.000 |
| :ATTenuation? | Query attenuation set-point |  |  | 7.0.000 |
| :DATa<i> | Path to set up/query calibration data: |  |  | 7.0.000 |
| :ATTenuation <f> | Specify attenuation |  |  | 7.0.000 |
| :ATTenuation? | Query attenuation |  |  | 7.0.000 |
| :VOLTage <f> | Specify voltage | V |  | 7.0.000 |
| :VOLTage? | Query voltage | V |  | 7.0.000 |
| :MODulation | Path to set up/query modulation settings: |  |  | 7.0.000 |
| :AMPLitude <f> | Specify amplitude | V |  | 7.0.000 |
| :AMPLitude? | Query amplitude | V |  | 7.0.000 |
| :FREQuency <f> | Specify frequency | Hz |  | 7.0.000 |
| :FREQuency? | Query frequency | Hz |  | 7.0.000 |
| :STATe <b> | Specify modulation state (0 = disabled, 1 = enabled) |  |  | 7.0.000 |
| :STATe? | Query modulation state |  |  | 7.0.000 |
| :WAVeform <i> | Specify waveform |  |  | 7.0.000 |
| :WAVeform? | Query waveform |  |  | 7.0.000 |
| :VOLTage <f> | Specify voltage set-point | V |  | 7.0.000 |
| :VOLTage? | Query voltage set-point | V |  | 7.0.000 |
| :VMONitor? | Query monitor voltage | V |  | 7.0.000 |
| :VREFerence<1¦2>? | Query reference voltage | bits |  | 6.0.000 |

Table 5.33 – SYSTem Command Summary (continued)

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :WADJust | Path to set up/query wavelength adjust settings: |  |  | Note 3 |
| :BAND <a> | Specify mode hop free band ID |  |  | Note 1 |
| :BAND? | Query mode hop free band ID |  |  | Note 1 |
| :FLAGs? | Query status flags |  |  | 4.9.002 |
| :INPut <i> | Specify input (0 = BNC socket, 1 = potentiometer) |  |  | 4.8.000 |
| :INPut? | Query input |  |  | 4.8.000 |
| :STATe <b> | Specify adjustment state (0 = disabled, 1 = enabled) |  |  | Note 2 |
| :STATe <i> | Specify adjustment state |  |  | 4.1.000 |
| :STATe? | Query adjustment state |  |  | Note 3 |
| :TOFFset? | Query temperature tuning offset | ºC |  | Note 1 |
| :WAVelength <f> | Specify wavelength | nm |  | 2.1.000 |
| :WAVelength? | Query wavelength | nm |  | 2.1.000 |
| :Shutdown | Path to set / query shutdown dwell |  |  | Note 4 |
| :DwellVal <i> | Specify dwell value | seconds |  | 8.5.008 |
| :DwellVal? | Query dwell value | seconds |  | 8.5.008 |
| :DState? | Query dwell state | boolean |  | 8.5.008 |
| :DMin <i> | Set minimum dwell | seconds |  | 8.5.008 |
| :DMin? | Query dwell minimum | seconds |  | 8.5.008 |
| :DMax <i> | Set dwell maximum | seconds |  | 8.5.008 |
| :DMax? | Query dwell maximum | seconds |  | 8.5.008 |

Notes

1. Supported by version 1.0.003a only.

2. Supported by version 1.0.003a and version 2.1.000 to 4.0.002 only.

3. Supported by version 1.0.003a and version 2.1.000 onwards.

4. Supported by version 8.5.008 and onwards

Table 5.34 – STATus Command Summary

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :STATus[0¦1¦2¦3] | Commands to query monitor readings and/or flags: |  |  | see text |
| :ALL? | Query monitor readings and flags |  |  | 5.0.000 |
| :FLAGs? | Query flags only |  |  | 5.0.000 |

Table 5.35 – MESSage Command Summary

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :MESSage <s> | Send message to remote device |  |  | 4.4.000 |
| :MESSage? | Read message from remote device |  |  | 4.4.000 |

Table 5.36 – DIAGnostic Command Summary

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :DIAGnostic | Commands to control diagnostic functions: |  |  | 5.2.000 |
| :AMPLifier | Path to set up/query amplifier diagnostics: |  |  | 5.2.000 |
| :ECHO <b> | Specify echo state (0 = disabled, 1 = enabled) |  |  | 5.2.000 |
| :ECHO? | Query echo state (0 = disabled, 1 = enabled) |  |  | 5.2.000 |
| :TEST | Path to set up/query test functions: |  |  | 5.3.004 |
| :EEPRom? | Query EEPROM test result |  |  | 7.2.000 |
| :LEDS <b> | Specify LED test state (0 = disabled, 1 = enabled) |  |  | 5.3.004 |
| :LEDS? | Query LED test state (0 = disabled, 1 = enabled) |  |  | 5.3.004 |
| :PADJust <b> | Specify power adjust test state (0 = disabled, 1 = enabled) |  |  | 5.3.004 |
| :PADJust? | Query power adjust test state (0 = disabled, 1 = enabled) |  |  | 5.3.004 |
| :VREFerence <b> | Specify Vref test state (0 = disabled, 1 = enabled) |  |  | 6.0.000 |
| :VREFerence? | Query Vref test state (0 = disabled, 1 = enabled) |  |  | 6.0.000 |
| :WDOG | Path to set up/query watchdog functions: |  |  | Note 1 |
| :BOOT | Reset and boot load via serial port |  |  | 6.0.000 |
| :FLAG <b> | Clear reset status flag (0 = no effect, 1 = clear flag) |  |  | Note 1 |
| :FLAG? | Query reset status flag (0 = external reset, 1 = w/dog reset) |  |  | Note 1 |
| :INTerrupt <b> | Specify interrupt enable state (0 = disabled, 1 = enabled) |  |  | Note 1 |
| :INTerrupt? | Query interrupt enable state (0 = disabled, 1 = enabled) |  |  | Note 1 |
| :STATe <b> | Specify watchdog enable state (0 = disabled, 1 = enabled) |  |  | Note 1 |
| :STATe? | Query watchdog enable state (0 = disabled, 1 = enabled) |  |  | Note 1 |
| :TEST | Test watchdog reset/interrupt |  |  | Note 1 |

Notes

1. Supported by versions 5.2.007, 5.3.005 and 5.4.002 onwards.

Table 5.37 – OPTimization Command Summary

| Command | Description | Units | Config. | Version |
| --- | --- | --- | --- | --- |
| :OPTimization | Commands to control optimization functions: |  |  | 5.4.000 |
| :DATA? | Query raw data |  |  | 5.4.000 |
| :FLAG? | Query confidence flag |  |  | 5.4.003 |
| :IDWell <i> | Specify initial dwell time | s/10 |  | Note 1 |
| :IDWell? | Query initial dwell time | s/10 |  | Note 1 |
| :LOG | Path to set up/query optimization log: |  |  | 6.5.000 |
| :CLEar | Clear optimization log |  |  | 6.5.000 |
| :EVENt<i> | Path to query event data: |  |  | 6.5.000 |
| :FLAG? | Query confidence flag |  |  | 6.5.000 |
| :TEMPerature? | Query heater temperature | °C |  | 6.5.000 |
| :TIME? | Query time of optimization | hours |  | 6.5.000 |
| :VOLTage? | Query optimized voltage | V |  | 6.5.000 |
| :INDex? | Query index of last event |  |  | 6.5.000 |
| :TOTal? | Query total number of events |  |  | 6.5.000 |
| :POINt<0¦1¦2>? | Query point data |  |  | 5.4.004 |
| :RAMP<0¦1> | Path to set up/query ramp data: |  |  | 5.4.003 |
| :DIPS? | Query dip positions |  |  | 5.4.003 |
| :PEAKs? | Query peak positions |  |  | 5.4.003 |
| :POWers? | Query power data |  |  | 5.4.007 |
| :QUADratic<0¦1>? | Query quadratic fit results |  |  | 5.4.003 |
| :SMLevels? | Query second mode level data |  |  | 5.4.007 |
| :SMRanges? | Query second mode range data |  |  | 6.3.000 |
| :STEPs? | Query step data |  |  | Note 3 |
| :RMAX <f> | Specify maximum ramp voltage | V |  | Note 2 |
| :RMAX? | Query maximum ramp voltage | V |  | Note 2 |
| :RMIN <f> | Specify minimum ramp voltage | V |  | Note 2 |
| :RMIN? | Query minimum ramp voltage | V |  | Note 2 |
| :RSTEPs <i> | Specify number of ramp steps |  |  | Note 1 |
| :RSTEPs? | Query number of ramp steps |  |  | Note 1 |
| :SDWell <i> | Specify step dwell time | s/10 |  | Note 1 |
| :SDWell? | Query step dwell time | s/10 |  | Note 1 |
| :SMAX <i> | Specify maximum step number |  |  | Note 4 |
| :SMAX? | Query maximum step number |  |  | Note 4 |
| :SMIN <i> | Specify minimum step number |  |  | Note 4 |
| :SMIN? | Query minimum step number |  |  | Note 4 |
| :STATe <b> | Specify state (0 = stop, 1 = start) |  |  | 5.4.000 |
| :STATe? | Query state (0 = stopped, 1 = running) |  |  | 5.4.000 |
| :VOPTimum? | Query optimum voltage | V |  | 5.4.003 |

Notes

1. Supported by versions 5.4.000 to 5.4.003 only.

2. Supported by versions 5.4.001 to 5.4.003 only.

3. Supported by versions 5.4.003 to 5.4.006 only.

4. Supported by version 5.4.000 only.

Table 5.38 – Error Messages

| Number | Description | Version |
| --- | --- | --- |
| -350 | Queue overflow | 1.0.000 |
| -320 | Storage fault | 1.0.000 |
| -314 | Save/recall memory lost | 1.0.000 |
| -224 | Illegal parameter value | 1.0.000 |
| -223 | Too much data | 1.0.000 |
| -222 | Data out of range | 1.0.000 |
| -221 | Settings conflict | 1.0.000 |
| -203 | Command protected | 1.0.000 |
| -154 | String too long | 1.0.000 |
| -151 | Invalid string data | 1.0.000 |
| -121 | Invalid character in number | 1.0.000 |
| -113 | Undefined header | 1.0.000 |
| -110 | Command header error | 1.0.000 |
| -109 | Missing parameter | 1.0.000 |
| -104 | Data type error | 1.0.000 |
| -102 | Syntax error | 1.0.000 |
| +000 | No error | 1.0.000 |
| +110 | Wavelength tuning error | 2.1.000 |
| +111 | No wavelength tuning data | 2.1.000 |
| +112 | Wavelength out of range | 2.1.000 |
| +113 | Wavelength identification error | 2.1.000 |
| +120 | Power adjust error | 4.8.000 |
| +122 | Power out of range | 4.8.000 |